## Sheet 1
Schema Title | Procedure Title | Question Title | Acceptance Criteria
Thermometer Calibration Send Out Annual | Annual Thermometer Calibration Send Out P1 | Date Sent Out Q1.P1 | Standard Acceptance Criteria Date Sent Out Q1.P1 Less than or equal to Today
Thermometer Calibration Send Out Annual | Annual Thermometer Calibration Send Out P1 | Calibrated Date Q4.P1 | Standard Acceptance Criteria Calibrated Date Q4.P1 Less than or equal to Today
Thermometer Calibration Send Out Annual | Annual Thermometer Calibration Send Out P1 | Problem with Equipment Q3.P1 | Standard Acceptance Criteria Problem with Equipment Q3.P1 Equal to No
CANTO As Needed | BD FACSCanto As Needed P1 | Regions Optimization Q1.P1 | Standard Acceptance Criteria Regions Optimization Q1.P1 Equal to Yes
CANTO As Needed | BD FACSCanto As Needed P1 | Problem with Equipment Q2.P1 | Standard Acceptance Criteria Problem with Equipment Q2.P1 Equal to No
Prism Start-up Daily | Prism Start-up Daily QC P1 | Load New Biohazard Waste Containers Q1.P1 | Standard Acceptance Criteria Load New Biohazard Waste Containers Q1.P1 Equal to Performed
Prism Start-up Daily | Prism Start-up Daily QC P1 | Prime Reagent Fluidics on All Assigned Channels Q2.P1 | Standard Acceptance Criteria Prime Reagent Fluidics on All Assigned Channels Q2.P1 Equal to Performed - OK
Prism Start-up Daily | Prism Start-up Daily QC P1 | Load New Reaction Trays Q3.P1 | Standard Acceptance Criteria Load New Reaction Trays Q3.P1 Equal to Performed
Prism Start-up Daily | Prism Start-up Daily QC P1 | Load New Pipette Tips Q4.P1 | Standard Acceptance Criteria Load New Pipette Tips Q4.P1 Equal to Performed
Prism Start-up Daily | Prism Start-up Daily QC P1 | Problem with Equipment Q5.P1 | Standard Acceptance Criteria Problem with Equipment Q5.P1 Equal to No
Prism As Needed | Prism As Needed QC P1 | Decontaminate Instrument Surfaces Q1.P1 | Standard Acceptance Criteria Decontaminate Instrument Surfaces Q1.P1 Equal to Performed
Prism As Needed | Prism As Needed QC P1 | Problem with Equipment Q2.P1 | Standard Acceptance Criteria Problem with Equipment Q2.P1 Equal to No
Prism Monthly | Prism Monthly QC P1 | Total Hours on Cleaner Treatment Q3.P1 | Standard Acceptance Criteria Total Hours on Cleaner Treatment Q3.P1 Between 2 to 64
Prism Monthly | Prism Monthly QC P1 | Purge Reagent Fluidics Solution Twice Q4.P1 | Standard Acceptance Criteria Purge Reagent Fluidics Solution Twice Q4.P1 Equal to Performed - OK
Prism Monthly | Prism Monthly QC P1 | Run Temperature Validation on Channel 1 Q5.P1 | Standard Acceptance Criteria Run Temperature Validation on Channel 1 Q5.P1 Equal to Passed
Prism Monthly | Prism Monthly QC P1 | Run Temperature Validation on Channel 2 Q6.P1 | Standard Acceptance Criteria Run Temperature Validation on Channel 2 Q6.P1 Equal to Passed
Prism Monthly | Prism Monthly QC P1 | Run Temperature Validation on Channel 3 Q7.P1 | Standard Acceptance Criteria Run Temperature Validation on Channel 3 Q7.P1 Equal to Passed
Prism Monthly | Prism Monthly QC P1 | Run Temperature Validation on Channel 4 Q8.P1 | Standard Acceptance Criteria Run Temperature Validation on Channel 4 Q8.P1 Equal to Passed
Prism Monthly | Prism Monthly QC P1 | Run Temperature Validation on Channel 5 Q9.P1 | Standard Acceptance Criteria Run Temperature Validation on Channel 5 Q9.P1 Equal to Passed
Prism Monthly | Prism Monthly QC P1 | Run Temperature Validation on Channel 6 Q10.P1 | Standard Acceptance Criteria Run Temperature Validation on Channel 6 Q10.P1 Equal to Passed
Prism Monthly | Prism Monthly QC P1 | Clean Optics Cover Filter Q11.P1 | Standard Acceptance Criteria Clean Optics Cover Filter Q11.P1 Equal to Performed
Prism Monthly | Prism Monthly QC P1 | Perform Optics Module Validation on Channel 1 Q12.P1 | Standard Acceptance Criteria Perform Optics Module Validation on Channel 1 Q12.P1 Equal to Passed
Prism Monthly | Prism Monthly QC P1 | Perform Optics Module Validation on Channel 2 Q13.P1 | Standard Acceptance Criteria Perform Optics Module Validation on Channel 2 Q13.P1 Equal to Passed
Prism Monthly | Prism Monthly QC P1 | Perform Optics Module Validation on Channel 3 Q14.P1 | Standard Acceptance Criteria Perform Optics Module Validation on Channel 3 Q14.P1 Equal to Passed
Prism Monthly | Prism Monthly QC P1 | Perform Optics Module Validation on Channel 4 Q15.P1 | Standard Acceptance Criteria Perform Optics Module Validation on Channel 4 Q15.P1 Equal to Passed
Prism Monthly | Prism Monthly QC P1 | Perform Optics Module Validation on Channel 5 Q16.P1 | Standard Acceptance Criteria Perform Optics Module Validation on Channel 5 Q16.P1 Equal to Passed
Prism Monthly | Prism Monthly QC P1 | Perform Optics Module Validation on Channel 6 Q17.P1 | Standard Acceptance Criteria Perform Optics Module Validation on Channel 6 Q17.P1 Equal to Passed
Prism Monthly | Prism Monthly QC P1 | Problem with Equipment Q18.P1 | Standard Acceptance Criteria Problem with Equipment Q18.P1 Equal to No
Prism Quarterly | Prism Quarterly QC P1 | Clean Tray Loaders Q1.P1 | Standard Acceptance Criteria Clean Tray Loaders Q1.P1 Equal to Performed
Prism Quarterly | Prism Quarterly QC P1 | Clean Refrigerator External Filter Q2.P1 | Standard Acceptance Criteria Clean Refrigerator External Filter Q2.P1 Equal to Performed
Prism Quarterly | Prism Quarterly QC P1 | Problem with Equipment Q3.P1 | Standard Acceptance Criteria Problem with Equipment Q3.P1 Equal to No
Prism Weekly | QC Performed P1 | Was the QC Performed Q1.P1 | Standard Acceptance Criteria Was the QC Performed Q1.P1 Equal to No
Prism Weekly | Prism Weekly QC P1.P2(Con) | Disinfect Pipette Tip Chute and Funnel Q1.P1.P2(Con) | Standard Acceptance Criteria Disinfect Pipette Tip Chute and Funnel Q1.P1.P2(Con) Equal to Performed
Prism Weekly | Prism Weekly QC P1.P2(Con) | Reset Log History Q2.P1.P2(Con) | Standard Acceptance Criteria Reset Log History Q2.P1.P2(Con) Equal to Performed
Prism Weekly | Prism Weekly QC P1.P2(Con) | Reset Event Log Q3.P1.P2(Con) | Standard Acceptance Criteria Reset Event Log Q3.P1.P2(Con) Equal to Performed
Prism Weekly | Prism Weekly QC P1.P2(Con) | Problem with Equipment Q4.P1.P2(Con) | Standard Acceptance Criteria Problem with Equipment Q4.P1.P2(Con) Equal to No
CANTO Bi-Annual | BD FACSCanto Bi-Annual P1 | Clean or Replace Air Filter Q1.P1 | Standard Acceptance Criteria Clean or Replace Air Filter Q1.P1 Equal to Performed
CANTO Bi-Annual | BD FACSCanto Bi-Annual P1 | Replace Fluidics Filters Q2.P1 | Standard Acceptance Criteria Replace Fluidics Filters Q2.P1 Equal to Performed
CANTO Bi-Annual | BD FACSCanto Bi-Annual P1 | Problem with Equipment Q3.P1 | Standard Acceptance Criteria Problem with Equipment Q3.P1 Equal to No
Coulter AcT 2 Diff Start-up Daily | Coulter Lot P1 | Normal Control Open Bottle Expiration Date Q1.P1 | Standard Acceptance Criteria Normal Control Open Bottle Expiration Date Q1.P1 Greater than or equal to Today
Coulter AcT 2 Diff Start-up Daily | Coulter Lot P1 | High Control Open Bottle Expiration Date Q2.P1 | Standard Acceptance Criteria High Control Open Bottle Expiration Date Q2.P1 Greater than or equal to Today
Coulter AcT 2 Diff Start-up Daily | Startup Check P2 | Start-Up Check Q1.P2 | Standard Acceptance Criteria Start-Up Check Q1.P2 Equal to Performed - OK
Coulter AcT 2 Diff Start-up Daily | Coulter Daily P3 | Normal Control Value-WBC Q1.P3 | Standard Acceptance Criteria Normal Control Value-WBC Q1.P3 Between 9.3 to 10.2
Coulter AcT 2 Diff Start-up Daily | Coulter Daily P3 | Normal Control Value - RBC Q2.P3 | Standard Acceptance Criteria Normal Control Value - RBC Q2.P3 Between 3.84 to 4.2
Coulter AcT 2 Diff Start-up Daily | Coulter Daily P3 | Normal Control Value - HGB Q3.P3 | Standard Acceptance Criteria Normal Control Value - HGB Q3.P3 Between 11 to 12.2
Coulter AcT 2 Diff Start-up Daily | Coulter Daily P3 | Normal Control Value - HCT Q4.P3 | Standard Acceptance Criteria Normal Control Value - HCT Q4.P3 Between 31.2 to 34.8
Coulter AcT 2 Diff Start-up Daily | Coulter Daily P3 | Normal Control Value PLT Q5.P3 | Standard Acceptance Criteria Normal Control Value PLT Q5.P3 Between 176 to 254
Coulter AcT 2 Diff Start-up Daily | Coulter Daily P3 | Abnormal High Control Value - WBC Q6.P3 | Standard Acceptance Criteria Abnormal High Control Value - WBC Q6.P3 Between 17.8 to 19.6
Coulter AcT 2 Diff Start-up Daily | Coulter Daily P3 | Abnormal High Control Value - RBC Q7.P3 | Standard Acceptance Criteria Abnormal High Control Value - RBC Q7.P3 Between 4.97 to 5.45
Coulter AcT 2 Diff Start-up Daily | Coulter Daily P3 | Abnormal High Control Value - HGB Q8.P3 | Standard Acceptance Criteria Abnormal High Control Value - HGB Q8.P3 Between 16.1 to 17.3
Coulter AcT 2 Diff Start-up Daily | Coulter Daily P3 | Abnormal High Control Value - HCT Q9.P3 | Standard Acceptance Criteria Abnormal High Control Value - HCT Q9.P3 Between 44.9 to 49.7
Coulter AcT 2 Diff Start-up Daily | Coulter Daily P3 | Abnormal High Control Value - PLT Q10.P3 | Standard Acceptance Criteria Abnormal High Control Value - PLT Q10.P3 Between 319 to 427
Coulter AcT 2 Diff Start-up Daily | Coulter Daily P3 | Problem with Equipment Q11.P3 | Standard Acceptance Criteria Problem with Equipment Q11.P3 Equal to No
Trima PM Bi-Annual | PM P1 | PM Performed Q3.P1 | Standard Acceptance Criteria PM Performed Q3.P1 Equal to Performed - OK
Trima PM Bi-Annual | PM P1 | Problem with Equipment Q4.P1 | Standard Acceptance Criteria Problem with Equipment Q4.P1 Equal to No
Trima Weekly | Trima Weekly P1 | System Surface Disinfected Q1.P1 | Standard Acceptance Criteria System Surface Disinfected Q1.P1 Equal to Performed - OK
Trima Weekly | Trima Weekly P1 | Sensors and Valves Cleaned Q2.P1 | Standard Acceptance Criteria Sensors and Valves Cleaned Q2.P1 Equal to Performed - OK
Trima Weekly | Trima Weekly P1 | Problem with Equipment Q3.P1 | Standard Acceptance Criteria Problem with Equipment Q3.P1 Equal to No
Trima Monthly | Trima Monthly P1 | Pump Housing and Rotors Cleaned Q1.P1 | Standard Acceptance Criteria Pump Housing and Rotors Cleaned Q1.P1 Equal to Performed - OK
Trima Monthly | Trima Monthly P1 | Fluid Leak Detector Cleaned Q2.P1 | Standard Acceptance Criteria Fluid Leak Detector Cleaned Q2.P1 Equal to Performed - OK
Trima Monthly | Trima Monthly P1 | Filter Removed and Cleaned Q3.P1 | Standard Acceptance Criteria Filter Removed and Cleaned Q3.P1 Equal to Performed - OK
Trima Monthly | Trima Monthly P1 | Problem with Equipment Q4.P1 | Standard Acceptance Criteria Problem with Equipment Q4.P1 Equal to No
Sealer Weekly | Sealer Weekly P1 | Clean/Wipe the Heat Sealing Element Q1.P1 | Standard Acceptance Criteria Clean/Wipe the Heat Sealing Element Q1.P1 Equal to Performed - OK
Sealer Weekly | Sealer Weekly P1 | Sealer Disinfected Q2.P1 | Standard Acceptance Criteria Sealer Disinfected Q2.P1 Equal to Performed
Sealer Weekly | Sealer Weekly P1 | Problem with Equipment Q3.P1 | Standard Acceptance Criteria Problem with Equipment Q3.P1 Equal to No
Vital Sign Monitor PM Annual | PM P1 | PM Performed Q3.P1 | Standard Acceptance Criteria PM Performed Q3.P1 Equal to Performed - OK
Vital Sign Monitor PM Annual | PM P1 | Problem with Equipment Q4.P1 | Standard Acceptance Criteria Problem with Equipment Q4.P1 Equal to No
Cell Washer Daily | Cell Washer Daily P1 | Clean the Cell Washer Q1.P1 | Standard Acceptance Criteria Clean the Cell Washer Q1.P1 Equal to Performed - OK
Cell Washer Daily | Cell Washer Daily P1 | Check Connections Q2.P1 | Standard Acceptance Criteria Check Connections Q2.P1 Equal to Performed - OK
Cell Washer Daily | Cell Washer Daily P1 | Record Saline Volume Q3.P1 | Standard Acceptance Criteria Record Saline Volume Q3.P1 Between 39 to 41
Cell Washer Daily | Cell Washer Daily P1 | Problem with Equipment Q4.P1 | Standard Acceptance Criteria Problem with Equipment Q4.P1 Equal to No
Electronic Balance Daily | Electronic Balance 5g P1.P2(Con) | Electronic Balance 5g Q1.P1.P2(Con) | Standard Acceptance Criteria Electronic Balance 5g Q1.P1.P2(Con) Between 4.99 to 5.01
Electronic Balance Daily | Electronic Balance 5g P1.P2(Con) | Problem with Equipment Q2.P1.P2(Con) | Standard Acceptance Criteria Problem with Equipment Q2.P1.P2(Con) Equal to No
Electronic Balance Daily | Electronic Balance 200g P1.P3(Con) | Electronic Balance 200g Q2.P1.P3(Con) | Standard Acceptance Criteria Electronic Balance 200g Q2.P1.P3(Con) Between 199.9 to 200.1
Electronic Balance Daily | Electronic Balance 200g P1.P3(Con) | Problem with Equipment Q1.P1.P3(Con) | Standard Acceptance Criteria Problem with Equipment Q1.P1.P3(Con) Equal to No
Electronic Balance Daily | Electronic Balance 300g P1.P4(Con) | Electronic Balance 300g Q1.P1.P4(Con) | Standard Acceptance Criteria Electronic Balance 300g Q1.P1.P4(Con) Between 299.9 to 300.1
Electronic Balance Daily | Electronic Balance 300g P1.P4(Con) | Problem with Equipment Q2.P1.P4(Con) | Standard Acceptance Criteria Problem with Equipment Q2.P1.P4(Con) Equal to No
Coulter AcT 2 Diff Hematology Analyzer Weekly | Coulter Weekly P1 | Clean Transducer Q1.P1 | Standard Acceptance Criteria Clean Transducer Q1.P1 Equal to Performed
Coulter AcT 2 Diff Hematology Analyzer Weekly | Coulter Weekly P1 | Clean Sample Holders Q2.P1 | Standard Acceptance Criteria Clean Sample Holders Q2.P1 Equal to Performed
Coulter AcT 2 Diff Hematology Analyzer Weekly | Coulter Weekly P1 | Problem with Equipment Q3.P1 | Standard Acceptance Criteria Problem with Equipment Q3.P1 Equal to No
CANTO Monthly | BD FACSCanto Monthly P1 | Archive Data Q5.P1 | Standard Acceptance Criteria Archive Data Q5.P1 Equal to Performed
CANTO Monthly | BD FACSCanto Monthly P1 | Long Clean Q1.P1 | Standard Acceptance Criteria Long Clean Q1.P1 Equal to Performed
CANTO Monthly | BD FACSCanto Monthly P1 | Change Waste Container Cap Q2.P1 | Standard Acceptance Criteria Change Waste Container Cap Q2.P1 Equal to Performed
CANTO Monthly | BD FACSCanto Monthly P1 | Print Levy-Jennings Report Q3.P1 | Standard Acceptance Criteria Print Levy-Jennings Report Q3.P1 Equal to Performed
CANTO Monthly | BD FACSCanto Monthly P1 | Problem with Equipment Q4.P1 | Standard Acceptance Criteria Problem with Equipment Q4.P1 Equal to No
CANTO Weekly | BD FACSCanto Weekly P1 | Clean Instrument with 10% Bleach Q1.P1 | Standard Acceptance Criteria Clean Instrument with 10% Bleach Q1.P1 Equal to Performed
CANTO Weekly | BD FACSCanto Weekly P1 | Purging Fluidics Filters Q2.P1 | Standard Acceptance Criteria Purging Fluidics Filters Q2.P1 Equal to Performed
CANTO Weekly | BD FACSCanto Weekly P1 | Empty Waste Tank Q3.P1 | Standard Acceptance Criteria Empty Waste Tank Q3.P1 Equal to Performed
CANTO Weekly | BD FACSCanto Weekly P1 | Clean Tube Sensor Q4.P1 | Standard Acceptance Criteria Clean Tube Sensor Q4.P1 Equal to Performed
CANTO Weekly | BD FACSCanto Weekly P1 | Problem with Equipment Q5.P1 | Standard Acceptance Criteria Problem with Equipment Q5.P1 Equal to No
Irradiator Daily | Problem with Equipment P1 | Problem with Equipment Q1.P1 | Standard Acceptance Criteria Problem with Equipment Q1.P1 Equal to No
Liquid Nitrogen Daily | Liquid Nitrogen Daily P1 | Level of Liquid Nitrogen Q1.P1 | Standard Acceptance Criteria Level of Liquid Nitrogen Q1.P1 Between 12 to 14
Liquid Nitrogen Daily | Liquid Nitrogen Added P1.P2(Con) | Liquid Nitrogen Added Q1.P1.P2(Con) | Standard Acceptance Criteria Liquid Nitrogen Added Q1.P1.P2(Con) Equal to Yes
Liquid Nitrogen Daily | Liquid Nitrogen Added P1.P2(Con) | Level of Liquid Nitrogen After Addition Q2.P1.P2(Con) | Standard Acceptance Criteria Level of Liquid Nitrogen After Addition Q2.P1.P2(Con) Between 12 to 14
Sorvall ST40 Centrifuge Daily | Sorvall ST40 Daily P1 | Clean Parts and Cabinets Q1.P1 | Standard Acceptance Criteria Clean Parts and Cabinets Q1.P1 Equal to Performed - OK
Sorvall ST40 Centrifuge Daily | Sorvall ST40 Daily P1 | Apply Corrosion Protection Oil Q2.P1 | Standard Acceptance Criteria Apply Corrosion Protection Oil Q2.P1 Equal to Performed - OK
Sorvall ST40 Centrifuge Daily | Sorvall ST40 Daily P1 | Apply Grease Q3.P1 | Standard Acceptance Criteria Apply Grease Q3.P1 Equal to Performed - OK
Sorvall ST40 Centrifuge Daily | Sorvall ST40 Daily P1 | Problem with Equipment Q4.P1 | Standard Acceptance Criteria Problem with Equipment Q4.P1 Equal to No
REES Daily | REES Daily P1 | Report Printed and Reviewed Q1.P1 | Standard Acceptance Criteria Report Printed and Reviewed Q1.P1 Equal to Yes
REES Daily | REES Daily P1 | Problem with Equipment Q2.P1 | Standard Acceptance Criteria Problem with Equipment Q2.P1 Equal to No
Sterile Connecting Device Weekly | Sterile Connecting Device (SCD) Weekly P1 | SCD Cleaned Q1.P1 | Standard Acceptance Criteria SCD Cleaned Q1.P1 Equal to Performed - OK
Sterile Connecting Device Weekly | Sterile Connecting Device (SCD) Weekly P1 | Problem with Equipment Q2.P1 | Standard Acceptance Criteria Problem with Equipment Q2.P1 Equal to No
Neo Daily | Immucor NEO Daily P1 | Software Shutdown Q1.P1 | Standard Acceptance Criteria Software Shutdown Q1.P1 Equal to Performed - OK
Neo Daily | Immucor NEO Daily P1 | Clean and Initialize Neo Q2.P1 | Standard Acceptance Criteria Clean and Initialize Neo Q2.P1 Equal to Performed - OK
Neo Daily | Immucor NEO Daily P1 | Check and Refill System Liquid Containers Q3.P1 | Standard Acceptance Criteria Check and Refill System Liquid Containers Q3.P1 Equal to Performed - OK
Neo Daily | Immucor NEO Daily P1 | Empty Waste Container Q4.P1 | Standard Acceptance Criteria Empty Waste Container Q4.P1 Equal to Performed - OK
Neo Daily | Immucor NEO Daily P1 | Check Pipettor Reference Q5.P1 | Standard Acceptance Criteria Check Pipettor Reference Q5.P1 Equal to Performed - OK
Neo Daily | Immucor NEO Daily P1 | Pipettor Self Check Q6.P1 | Standard Acceptance Criteria Pipettor Self Check Q6.P1 Equal to Performed - OK
Neo Daily | Immucor NEO Daily P1 | Verify Reader Q7.P1 | Standard Acceptance Criteria Verify Reader Q7.P1 Equal to Performed - OK
Neo Daily | Immucor NEO Daily P1 | Reagent QC Q9.P1 | Standard Acceptance Criteria Reagent QC Q9.P1 Equal to Acceptable
Neo Daily | Immucor NEO Daily P1 | Problem with Equipment Q10.P1 | Standard Acceptance Criteria Problem with Equipment Q10.P1 Equal to No
Neo Weekly | QC Performed P1 | Was the QC Performed Q1.P1 | Standard Acceptance Criteria Was the QC Performed Q1.P1 Equal to Yes
Neo Weekly | Immucor NEO Weekly P1.P2(Con) | Flow350: Each Channel Is Between 2-3 ml Q1.P1.P2(Con) | Standard Acceptance Criteria Flow350: Each Channel Is Between 2-3 ml Q1.P1.P2(Con) Equal to Yes
Neo Weekly | Immucor NEO Weekly P1.P2(Con) | Flow350: Channel to Channel Variation Is <0.5 ml Q2.P1.P2(Con) | Standard Acceptance Criteria Flow350: Channel to Channel Variation Is <0.5 ml Q2.P1.P2(Con) Equal to Yes
Neo Weekly | Immucor NEO Weekly P1.P2(Con) | Flow350: Record Estimated Average Volume Q3.P1.P2(Con) | Standard Acceptance Criteria Flow350: Record Estimated Average Volume Q3.P1.P2(Con) Between 2 to 3
Neo Weekly | Immucor NEO Weekly P1.P2(Con) | Flow200: Each Channel Is Between 2-3 ml Q4.P1.P2(Con) | Standard Acceptance Criteria Flow200: Each Channel Is Between 2-3 ml Q4.P1.P2(Con) Equal to Yes
Neo Weekly | Immucor NEO Weekly P1.P2(Con) | Flow200: Channel to Channel Variation Is <0.5 ml Q5.P1.P2(Con) | Standard Acceptance Criteria Flow200: Channel to Channel Variation Is <0.5 ml Q5.P1.P2(Con) Equal to Yes
Neo Weekly | Immucor NEO Weekly P1.P2(Con) | Flow200: Record Estimated Average Volume Q6.P1.P2(Con) | Standard Acceptance Criteria Flow200: Record Estimated Average Volume Q6.P1.P2(Con) Between 2 to 3
Neo Weekly | Immucor NEO Weekly P1.P2(Con) | Reader Performance Q7.P1.P2(Con) | Standard Acceptance Criteria Reader Performance Q7.P1.P2(Con) Equal to Performed - OK
Neo Weekly | Immucor NEO Weekly P1.P2(Con) | Add RelyOn to the Pipettor Wash Towers Q8.P1.P2(Con) | Standard Acceptance Criteria Add RelyOn to the Pipettor Wash Towers Q8.P1.P2(Con) Equal to Performed - OK
Neo Weekly | Immucor NEO Weekly P1.P2(Con) | Empty and Rinse Common Waste Container with RelyOn Q9.P1.P2(Con) | Standard Acceptance Criteria Empty and Rinse Common Waste Container with RelyOn Q9.P1.P2(Con) Equal to Performed - OK
Neo Weekly | Immucor NEO Weekly P1.P2(Con) | Empty and Rinse Shuttle Container with RelyOn Q10.P1.P2(Con) | Standard Acceptance Criteria Empty and Rinse Shuttle Container with RelyOn Q10.P1.P2(Con) Equal to Performed - OK
Neo Weekly | Immucor NEO Weekly P1.P2(Con) | Rinse the Common Waste Container with DI H2O Q11.P1.P2(Con) | Standard Acceptance Criteria Rinse the Common Waste Container with DI H2O Q11.P1.P2(Con) Equal to Performed - OK
Neo Weekly | Immucor NEO Weekly P1.P2(Con) | Rinse the Shuttle Container with DI H2O Q12.P1.P2(Con) | Standard Acceptance Criteria Rinse the Shuttle Container with DI H2O Q12.P1.P2(Con) Equal to Performed - OK
Neo Weekly | Immucor NEO Weekly P1.P2(Con) | Clean Common Waste Container Lid with Isopropanol Q13.P1.P2(Con) | Standard Acceptance Criteria Clean Common Waste Container Lid with Isopropanol Q13.P1.P2(Con) Equal to Performed - OK
Neo Weekly | Immucor NEO Weekly P1.P2(Con) | Clean Shuttle Container Lid with Isopropanol Q14.P1.P2(Con) | Standard Acceptance Criteria Clean Shuttle Container Lid with Isopropanol Q14.P1.P2(Con) Equal to Performed - OK
Neo Weekly | Immucor NEO Weekly P1.P2(Con) | Inspect Syringes for Leakage and Tightness of Barrels Q15.P1.P2(Con) | Standard Acceptance Criteria Inspect Syringes for Leakage and Tightness of Barrels Q15.P1.P2(Con) Equal to Performed - OK
Neo Weekly | Immucor NEO Weekly P1.P2(Con) | Archive Q16.P1.P2(Con) | Standard Acceptance Criteria Archive Q16.P1.P2(Con) Equal to Performed - OK
Neo Weekly | Immucor NEO Weekly P1.P2(Con) | Problem with Equipment Q17.P1.P2(Con) | Standard Acceptance Criteria Problem with Equipment Q17.P1.P2(Con) Equal to No
Neo Monthly | Immucor NEO Monthly P1 | Clear Test Data Q1.P1 | Standard Acceptance Criteria Clear Test Data Q1.P1 Equal to Performed - OK
Neo Monthly | Immucor NEO Monthly P1 | Enter Reader Verification Plate Calibration Date Q2.P1 | Standard Acceptance Criteria Enter Reader Verification Plate Calibration Date Q2.P1 Less than or equal to Today
Neo Monthly | Immucor NEO Monthly P1 | Is This Date in the Last Three Years Q3.P1 | Standard Acceptance Criteria Is This Date in the Last Three Years Q3.P1 Equal to Yes
Neo Monthly | Immucor NEO Monthly P1 | Perform Decontamination Procedure for DB_CK Q4.P1 | Standard Acceptance Criteria Perform Decontamination Procedure for DB_CK Q4.P1 Equal to Performed - OK
Neo Monthly | Immucor NEO Monthly P1 | Perform Decontamination Procedure for PrimeDec Q5.P1 | Standard Acceptance Criteria Perform Decontamination Procedure for PrimeDec Q5.P1 Equal to Performed - OK
Neo Monthly | Immucor NEO Monthly P1 | Empty and Rinse Containers W/ RelyOn Followed by PBS Q6.P1 | Standard Acceptance Criteria Empty and Rinse Containers W/ RelyOn Followed by PBS Q6.P1 Equal to Performed - OK
Neo Monthly | Immucor NEO Monthly P1 | Perform Decontamination Procedure for ClearTube Q7.P1 | Standard Acceptance Criteria Perform Decontamination Procedure for ClearTube Q7.P1 Equal to Performed - OK
Neo Monthly | Immucor NEO Monthly P1 | Perform Decontamination for Prime All (Run 4 Times) Q8.P1 | Standard Acceptance Criteria Perform Decontamination for Prime All (Run 4 Times) Q8.P1 Equal to Performed - OK
Neo Monthly | Immucor NEO Monthly P1 | Refill with New PBS Q9.P1 | Standard Acceptance Criteria Refill with New PBS Q9.P1 Equal to Performed - OK
Neo Monthly | Immucor NEO Monthly P1 | Wipe System Liquid Sensors with Alcohol Q10.P1 | Standard Acceptance Criteria Wipe System Liquid Sensors with Alcohol Q10.P1 Equal to Performed - OK
Neo Monthly | Immucor NEO Monthly P1 | Perform Decontamination Procedure for HemCheck Q11.P1 | Standard Acceptance Criteria Perform Decontamination Procedure for HemCheck Q11.P1 Equal to Performed - OK
Neo Monthly | Immucor NEO Monthly P1 | Problem with Equipment Q12.P1 | Standard Acceptance Criteria Problem with Equipment Q12.P1 Equal to No
TempIR Non-Contact Thermometer Daily | TempIR P1 | TempIR Cleaning Q1.P1 | Standard Acceptance Criteria TempIR Cleaning Q1.P1 Equal to Performed - OK
TempIR Non-Contact Thermometer Daily | TempIR P1 | Start-Up Check Q2.P1 | Standard Acceptance Criteria Start-Up Check Q2.P1 Equal to Performed - OK
TempIR Non-Contact Thermometer Daily | TempIR P1 | Problem with Equipment Q4.P1 | Standard Acceptance Criteria Problem with Equipment Q4.P1 Equal to No
Hemoflow Daily | Hemoflow Daily Calibration P1 | Calibration Passed? Q1.P1 | Standard Acceptance Criteria Calibration Passed? Q1.P1 Equal to Passed
Hemoflow Daily | Hemoflow Daily Calibration P1 | Problem with Equipment Q2.P1 | Standard Acceptance Criteria Problem with Equipment Q2.P1 Equal to No
Irradiator PM Bi-Annual | Irradiator Bi-Annual P1 | Leak Test/Check Performed? Q2.P1 | Standard Acceptance Criteria Leak Test/Check Performed? Q2.P1 Equal to Performed - OK
Irradiator PM Bi-Annual | Irradiator Bi-Annual P1 | Electrical Check Performed? Q1.P1 | Standard Acceptance Criteria Electrical Check Performed? Q1.P1 Equal to Performed - OK
Irradiator PM Bi-Annual | Irradiator Bi-Annual P1 | Problem with Equipment Q3.P1 | Standard Acceptance Criteria Problem with Equipment Q3.P1 Equal to No
Sterile Connecting Device Daily | SCD Start-Up P1 | Start-Up Q1.P1 | Standard Acceptance Criteria Start-Up Q1.P1 Equal to Performed - OK
Hematron Weekly | Hematron Weekly P1 | Cleaned Q1.P1 | Standard Acceptance Criteria Cleaned Q1.P1 Equal to Performed - OK
Hematron Weekly | Hematron Weekly P1 | Problems with Equipment/QC Q2.P1 | Standard Acceptance Criteria Problems with Equipment/QC Q2.P1 Equal to No
IBL 437 Irradiator Weekly | IBL 437 Irradiator Weekly P1 | Clean and Disinfect Q1.P1 | Standard Acceptance Criteria Clean and Disinfect Q1.P1 Equal to Performed - OK
IBL 437 Irradiator Weekly | IBL 437 Irradiator Weekly P1 | Problems with Equipment/QC Q2.P1 | Standard Acceptance Criteria Problems with Equipment/QC Q2.P1 Equal to No
Gammacell 3000 Irradiator Quarterly | Gammacell 3000 Irradiator Quarterly P1 | Door Interlock Q1.P1 | Standard Acceptance Criteria Door Interlock Q1.P1 Equal to Passed
Gammacell 3000 Irradiator Quarterly | Gammacell 3000 Irradiator Quarterly P1 | Emergency Power Check Q2.P1 | Standard Acceptance Criteria Emergency Power Check Q2.P1 Equal to Passed
Gammacell 3000 Irradiator Quarterly | Gammacell 3000 Irradiator Quarterly P1 | Timer Check Q3.P1 | Standard Acceptance Criteria Timer Check Q3.P1 Equal to Passed
Gammacell 3000 Irradiator Quarterly | Gammacell 3000 Irradiator Quarterly P1 | Timer Check Q5.P1 | Standard Acceptance Criteria Timer Check Q5.P1 Equal to Passed
Gammacell 3000 Irradiator Quarterly | Gammacell 3000 Irradiator Quarterly P1 | Problem with Equipment Q4.P1 | Standard Acceptance Criteria Problem with Equipment Q4.P1 Equal to No
Refrigerated Centrifuge Weekly | Refrigerated Centrifuge Weekly P1 | Wash Rotor Chamber Q1.P1 | Standard Acceptance Criteria Wash Rotor Chamber Q1.P1 Equal to Performed
Refrigerated Centrifuge Weekly | Refrigerated Centrifuge Weekly P1 | Clean Cabinet Panels Q2.P1 | Standard Acceptance Criteria Clean Cabinet Panels Q2.P1 Equal to Performed
Refrigerated Centrifuge Weekly | Refrigerated Centrifuge Weekly P1 | Clean the Top Deck and Chamber Door Q3.P1 | Standard Acceptance Criteria Clean the Top Deck and Chamber Door Q3.P1 Equal to Performed
Refrigerated Centrifuge Weekly | Refrigerated Centrifuge Weekly P1 | Lubricate Cups Q4.P1 | Standard Acceptance Criteria Lubricate Cups Q4.P1 Equal to Performed
Refrigerated Centrifuge Weekly | Refrigerated Centrifuge Weekly P1 | Problem with Equipment Q5.P1 | Standard Acceptance Criteria Problem with Equipment Q5.P1 Equal to No
Freezer Check Weekly QC Duplicate | Freezer Check Weekly P1 | Brushed/Defrosted Freezer Q1.P1 | Standard Acceptance Criteria Brushed/Defrosted Freezer Q1.P1 Equal to Performed
Freezer Check Weekly QC Duplicate | Freezer Check Weekly P1 | Problem with Equipment Q2.P1 | Standard Acceptance Criteria Problem with Equipment Q2.P1 Equal to No
Eyewash Weekly | Eyewash Weekly P1 | Check Area Is Unobstructed Q1.P1 | Standard Acceptance Criteria Check Area Is Unobstructed Q1.P1 Equal to Performed - OK
Eyewash Weekly | Eyewash Weekly P1 | Covers Present and Function Appropriately Q2.P1 | Standard Acceptance Criteria Covers Present and Function Appropriately Q2.P1 Equal to Performed - OK
Eyewash Weekly | Eyewash Weekly P1 | Drainage Acceptable Q3.P1 | Standard Acceptance Criteria Drainage Acceptable Q3.P1 Equal to Performed - OK
Eyewash Weekly | Eyewash Weekly P1 | Activation Performed and Acceptable Q4.P1 | Standard Acceptance Criteria Activation Performed and Acceptable Q4.P1 Equal to Performed - OK
Eyewash Weekly | Eyewash Weekly P1 | Problem with Equipment Q5.P1 | Standard Acceptance Criteria Problem with Equipment Q5.P1 Equal to No
Eyewash Molecular Area Weekly-inactive | Eyewash Weekly Molecular Area P1 | Eyewash Location Q1.P1 | Standard Acceptance Criteria Eyewash Location Q1.P1 Equal to Molecular
Eyewash Molecular Area Weekly-inactive | Eyewash Weekly Molecular Area P1 | Check Area Is Unobstructed? Q2.P1 | Standard Acceptance Criteria Check Area Is Unobstructed? Q2.P1 Equal to Performed - OK
Eyewash Molecular Area Weekly-inactive | Eyewash Weekly Molecular Area P1 | Covers Present and Function Appropriately? Q3.P1 | Standard Acceptance Criteria Covers Present and Function Appropriately? Q3.P1 Equal to Performed - OK
Eyewash Molecular Area Weekly-inactive | Eyewash Weekly Molecular Area P1 | Activation Q4.P1 | Standard Acceptance Criteria Activation Q4.P1 Equal to Performed - OK
Eyewash Molecular Area Weekly-inactive | Eyewash Weekly Molecular Area P1 | Drainage Acceptable? Q5.P1 | Standard Acceptance Criteria Drainage Acceptable? Q5.P1 Equal to Performed - OK
Eyewash Molecular Area Weekly-inactive | Eyewash Weekly Molecular Area P1 | Problem with Equipment Q6.P1 | Standard Acceptance Criteria Problem with Equipment Q6.P1 Equal to No
CMV New Lot Qualification As Needed | CMV New Lot Qualification P3 | Current Plate + Negative Internal Control Q1.P3 | Standard Acceptance Criteria Current Plate + Negative Internal Control Q1.P3 Equal to Negative:Acceptable
CMV New Lot Qualification As Needed | CMV New Lot Qualification P3 | Current Plate + Positive Internal Control Q2.P3 | Standard Acceptance Criteria Current Plate + Positive Internal Control Q2.P3 Equal to Positive:Acceptable
CMV New Lot Qualification As Needed | CMV New Lot Qualification P3 | Current Plate + Negative External Control Q3.P3 | Standard Acceptance Criteria Current Plate + Negative External Control Q3.P3 Equal to Negative - Acceptable
CMV New Lot Qualification As Needed | CMV New Lot Qualification P3 | Current Plate + Positive External Control Q4.P3 | Standard Acceptance Criteria Current Plate + Positive External Control Q4.P3 Equal to Positive - Acceptable
CMV New Lot Qualification As Needed | CMV New Lot Qualification P3 | Current Plate + Known Negative Patient Q5.P3 | Standard Acceptance Criteria Current Plate + Known Negative Patient Q5.P3 Equal to Negative - Acceptable
CMV New Lot Qualification As Needed | CMV New Lot Qualification P3 | Current Plate + Known Positive Patient Q6.P3 | Standard Acceptance Criteria Current Plate + Known Positive Patient Q6.P3 Equal to Positive - Acceptable
CMV New Lot Qualification As Needed | CMV New Lot Qualification P3 | New Plate + Negative Internal Control Q7.P3 | Standard Acceptance Criteria New Plate + Negative Internal Control Q7.P3 Equal to Negative:Acceptable
CMV New Lot Qualification As Needed | CMV New Lot Qualification P3 | New Plate + Positive Internal Control Q8.P3 | Standard Acceptance Criteria New Plate + Positive Internal Control Q8.P3 Equal to Positive:Acceptable
CMV New Lot Qualification As Needed | CMV New Lot Qualification P3 | New Plate + Negative External Control Q9.P3 | Standard Acceptance Criteria New Plate + Negative External Control Q9.P3 Equal to Negative - Acceptable
CMV New Lot Qualification As Needed | CMV New Lot Qualification P3 | New Plate + Positive External Control Q10.P3 | Standard Acceptance Criteria New Plate + Positive External Control Q10.P3 Equal to Positive - Acceptable
CMV New Lot Qualification As Needed | CMV New Lot Qualification P3 | New Plate + Known Negative Patient Q11.P3 | Standard Acceptance Criteria New Plate + Known Negative Patient Q11.P3 Equal to Negative - Acceptable
CMV New Lot Qualification As Needed | CMV New Lot Qualification P3 | New Plate + Known Positive Patient Q12.P3 | Standard Acceptance Criteria New Plate + Known Positive Patient Q12.P3 Equal to Positive - Acceptable
CMV New Lot Qualification As Needed | CMV New Lot Qualification P3 | CMV New Lot Qualification Acceptable Q13.P3 | Standard Acceptance Criteria CMV New Lot Qualification Acceptable Q13.P3 Equal to Acceptable
CMV New Lot Qualification As Needed | CMV New Lot Qualification P3 | Problem with Equipment Q14.P3 | Standard Acceptance Criteria Problem with Equipment Q14.P3 Equal to No
Hemocue Start-up Daily | Hemocue Start-up Daily QC P6 | Normal Control Open Bottle Expiration Date Q5.P6 | Standard Acceptance Criteria Normal Control Open Bottle Expiration Date Q5.P6 Greater than or equal to Today
Hemocue Start-up Daily | Hemocue Start-up Daily QC P6 | High Control Open Bottle Expiration Date Q6.P6 | Standard Acceptance Criteria High Control Open Bottle Expiration Date Q6.P6 Greater than or equal to Today
Hemocue Start-up Daily | Hemocue Start-up Daily QC P6 | Cuvette Open Bottle Expiration Date Q7.P6 | Standard Acceptance Criteria Cuvette Open Bottle Expiration Date Q7.P6 Greater than or equal to Today
Hemocue Start-up Daily | Hemocue Start-up Daily QC P6 | Start-Up Check Q3.P6 | Standard Acceptance Criteria Start-Up Check Q3.P6 Equal to Performed - OK
Hemocue Start-up Daily | Hemocue Start-up Daily QC P6 | Normal Control Reading Q1.P6 | Standard Acceptance Criteria Normal Control Reading Q1.P6 Greater than 0
Hemocue Start-up Daily | Hemocue Start-up Daily QC P6 | High Control Reading Q2.P6 | Standard Acceptance Criteria High Control Reading Q2.P6 Greater than 0
Hemocue Start-up Daily | Hemocue Start-up Daily QC P6 | Problem with Equipment Q4.P6 | Standard Acceptance Criteria Problem with Equipment Q4.P6 Equal to No
Cell Washer Monthly | Cell Washer Monthly P1 | Decontamination Q1.P1 | Standard Acceptance Criteria Decontamination Q1.P1 Equal to Performed - OK
Cell Washer Monthly | Cell Washer Monthly P1 | Problem with Equipment Q2.P1 | Standard Acceptance Criteria Problem with Equipment Q2.P1 Equal to No
TS Electronic Balance Monthly | Electronic Balance Monthly P1 | Bubble Centered Q2.P1 | Standard Acceptance Criteria Bubble Centered Q2.P1 Equal to Yes
TS Electronic Balance Monthly | Electronic Balance Monthly P1 | Readout (g) After Tared Q3.P1 | Standard Acceptance Criteria Readout (g) After Tared Q3.P1 Between -0.1 to 0.1
TS Electronic Balance Monthly | Electronic Balance Monthly TS Standard P1.P2(Con) | Readout 5 G Weight Q1.P1.P2(Con) | Standard Acceptance Criteria Readout 5 G Weight Q1.P1.P2(Con) Between 4.9 to 5.1
TS Electronic Balance Monthly | Electronic Balance Monthly TS Standard P1.P2(Con) | Readout 30 G Weight Q2.P1.P2(Con) | Standard Acceptance Criteria Readout 30 G Weight Q2.P1.P2(Con) Between 29.9 to 30.1
TS Electronic Balance Monthly | Electronic Balance Monthly TS Standard P1.P2(Con) | Readout 100 G Weight Q3.P1.P2(Con) | Standard Acceptance Criteria Readout 100 G Weight Q3.P1.P2(Con) Between 99.9 to 100.1
TS Electronic Balance Monthly | Electronic Balance Monthly TS Standard P1.P2(Con) | Problem with Equipment Q4.P1.P2(Con) | Standard Acceptance Criteria Problem with Equipment Q4.P1.P2(Con) Equal to No
TS Electronic Balance Monthly | Electronic Balance Monthly NEO P1.P2(Con) | Readout 1 G Weight Q1.P1.P2(Con) | Standard Acceptance Criteria Readout 1 G Weight Q1.P1.P2(Con) Between 0.99 to 1.01
TS Electronic Balance Monthly | Electronic Balance Monthly NEO P1.P2(Con) | Readout 2 G Weight Q2.P1.P2(Con) | Standard Acceptance Criteria Readout 2 G Weight Q2.P1.P2(Con) Between 1.99 to 2.01
TS Electronic Balance Monthly | Electronic Balance Monthly NEO P1.P2(Con) | Readout 3 G Weight Q3.P1.P2(Con) | Standard Acceptance Criteria Readout 3 G Weight Q3.P1.P2(Con) Between 2.99 to 3.01
TS Electronic Balance Monthly | Electronic Balance Monthly NEO P1.P2(Con) | Readout 30 G Weight Q4.P1.P2(Con) | Standard Acceptance Criteria Readout 30 G Weight Q4.P1.P2(Con) Between 29.9 to 30.1
TS Electronic Balance Monthly | Electronic Balance Monthly NEO P1.P2(Con) | Problem with Equipment Q5.P1.P2(Con) | Standard Acceptance Criteria Problem with Equipment Q5.P1.P2(Con) Equal to No
Irradiator Quarterly-inactive | Irradiator Quarterly QC P4 | Timer Check Q1.P4 | Standard Acceptance Criteria Timer Check Q1.P4 Equal to Passed
Irradiator Quarterly-inactive | Irradiator Quarterly QC P4 | Door Interlock Q2.P4 | Standard Acceptance Criteria Door Interlock Q2.P4 Equal to Passed
Irradiator Quarterly-inactive | Irradiator Quarterly QC P4 | Emergency Power Check Q3.P4 | Standard Acceptance Criteria Emergency Power Check Q3.P4 Equal to Passed
Irradiator Quarterly-inactive | Irradiator Quarterly QC P4 | Problem with Equipment Q4.P4 | Standard Acceptance Criteria Problem with Equipment Q4.P4 Equal to No
Heatblock Daily | Heatblock Daily P1 | Record Heatblock Incubator Temperature Q1.P1 | Standard Acceptance Criteria Record Heatblock Incubator Temperature Q1.P1 Between 36 to 38
Sorvall ST40 Centrifuge Monthly | Sorvall ST40 Centrifuge Monthly P1 | Disinfect Rotor and Parts Q1.P1 | Standard Acceptance Criteria Disinfect Rotor and Parts Q1.P1 Equal to Performed - OK
Sorvall ST40 Centrifuge Monthly | Sorvall ST40 Centrifuge Monthly P1 | Apply Grease Q2.P1 | Standard Acceptance Criteria Apply Grease Q2.P1 Equal to Performed - OK
Sorvall ST40 Centrifuge Monthly | Sorvall ST40 Centrifuge Monthly P1 | Problem with Equipment Q3.P1 | Standard Acceptance Criteria Problem with Equipment Q3.P1 Equal to No
DI Water Dispenser Monthly | DI Water Monthly P1 | Type II Sterility Results Q1.P1 | Standard Acceptance Criteria Type II Sterility Results Q1.P1 Between 0 to 100
DI Water Dispenser Monthly | DI Water Monthly P1 | Problem with Equipment Q2.P1 | Standard Acceptance Criteria Problem with Equipment Q2.P1 Equal to No
Sorvall RC3BP+ Daily | Sorvall RC3BP+ Daily P1 | QC Run Q1.P1 | Standard Acceptance Criteria QC Run Q1.P1 Equal to Passed
Sorvall RC3BP+ Daily | Sorvall RC3BP+ Daily P1 | Problem with Equipment Q2.P1 | Standard Acceptance Criteria Problem with Equipment Q2.P1 Equal to No
Refrigerated Centrifuge Bi-Annual | Refrigerated Centrifuge Bi-Annual P1 | Record Centrifuge RPM Setting 1st Spin Q1.P1 | Standard Acceptance Criteria Record Centrifuge RPM Setting 1st Spin Q1.P1 Between 1 to 9999
Refrigerated Centrifuge Bi-Annual | Refrigerated Centrifuge Bi-Annual P1 | Record Centrifuge RPM Actual Spin Q2.P1 | Standard Acceptance Criteria Record Centrifuge RPM Actual Spin Q2.P1 Between 1 to 9999
Refrigerated Centrifuge Bi-Annual | Refrigerated Centrifuge Bi-Annual P1 | Actual Spin RPM ± 100 of Centrifuge Setting Q3.P1 | Standard Acceptance Criteria Actual Spin RPM ± 100 of Centrifuge Setting Q3.P1 Equal to Yes
Refrigerated Centrifuge Bi-Annual | Refrigerated Centrifuge Bi-Annual P1 | Record Centrifuge RPM Setting 2nd Spin Q4.P1 | Standard Acceptance Criteria Record Centrifuge RPM Setting 2nd Spin Q4.P1 Between 1 to 9999
Refrigerated Centrifuge Bi-Annual | Refrigerated Centrifuge Bi-Annual P1 | Record Centrifuge RPM Actual Spin Q5.P1 | Standard Acceptance Criteria Record Centrifuge RPM Actual Spin Q5.P1 Between 1 to 9999
Refrigerated Centrifuge Bi-Annual | Refrigerated Centrifuge Bi-Annual P1 | Actual Spin RPM ± 100 of Centrifuge Setting Q6.P1 | Standard Acceptance Criteria Actual Spin RPM ± 100 of Centrifuge Setting Q6.P1 Equal to Yes
Refrigerated Centrifuge Bi-Annual | Refrigerated Centrifuge Bi-Annual P1 | Timer 5 Min. +/- 30 Sec Q7.P1 | Standard Acceptance Criteria Timer 5 Min. +/- 30 Sec Q7.P1 Less than 10:35:45
Refrigerated Centrifuge Bi-Annual | Refrigerated Centrifuge Bi-Annual P1 | Beginning Temperature 22°C Q8.P1 | Standard Acceptance Criteria Beginning Temperature 22°C Q8.P1 Between 20 to 24
Refrigerated Centrifuge Bi-Annual | Refrigerated Centrifuge Bi-Annual P1 | End Temperature 22°C Q9.P1 | Standard Acceptance Criteria End Temperature 22°C Q9.P1 Between 20 to 24
Refrigerated Centrifuge Bi-Annual | Refrigerated Centrifuge Bi-Annual P1 | Beginning Temperature 4°C Q10.P1 | Standard Acceptance Criteria Beginning Temperature 4°C Q10.P1 Between 2 to 6
Refrigerated Centrifuge Bi-Annual | Refrigerated Centrifuge Bi-Annual P1 | End Temperature 4°C Q11.P1 | Standard Acceptance Criteria End Temperature 4°C Q11.P1 Between 2 to 6
Refrigerated Centrifuge Bi-Annual | Refrigerated Centrifuge Bi-Annual P1 | Problem with Equipment Q12.P1 | Standard Acceptance Criteria Problem with Equipment Q12.P1 Equal to No
Tigris PM Bi-Annual duplicate | PM P1 | PM Performed Q3.P1 | Standard Acceptance Criteria PM Performed Q3.P1 Equal to Performed - OK
Tigris PM Bi-Annual duplicate | PM P1 | Problem with Equipment Q4.P1 | Standard Acceptance Criteria Problem with Equipment Q4.P1 Equal to No
Neo PM Bi-Annual | PM P1 | PM Performed Q3.P1 | Standard Acceptance Criteria PM Performed Q3.P1 Equal to Performed - OK
Neo PM Bi-Annual | PM P1 | Problem with Equipment Q4.P1 | Standard Acceptance Criteria Problem with Equipment Q4.P1 Equal to No
RPI PM Bi-Annual | PM P1 | PM Performed Q3.P1 | Standard Acceptance Criteria PM Performed Q3.P1 Equal to Performed - OK
RPI PM Bi-Annual | PM P1 | Problem with Equipment Q4.P1 | Standard Acceptance Criteria Problem with Equipment Q4.P1 Equal to No
IDT Electronic Balance Monthly | IDT Electronic Balance Monthly P1 | Readout (g) After Tared Q1.P1 | Standard Acceptance Criteria Readout (g) After Tared Q1.P1 Between -0.01 to 0.01
IDT Electronic Balance Monthly | IDT Electronic Balance Monthly P1 | Readout 1 G Weight Q2.P1 | Standard Acceptance Criteria Readout 1 G Weight Q2.P1 Between 0.99 to 1.01
IDT Electronic Balance Monthly | IDT Electronic Balance Monthly P1 | Readout 2 G Weight Q3.P1 | Standard Acceptance Criteria Readout 2 G Weight Q3.P1 Between 1.99 to 2.01
IDT Electronic Balance Monthly | IDT Electronic Balance Monthly P1 | Readout 3 G Weight Q4.P1 | Standard Acceptance Criteria Readout 3 G Weight Q4.P1 Between 2.99 to 3.01
IDT Electronic Balance Monthly | IDT Electronic Balance Monthly P1 | Readout 30 G Weight Q5.P1 | Standard Acceptance Criteria Readout 30 G Weight Q5.P1 Between 29.99 to 30.01
IDT Electronic Balance Monthly | IDT Electronic Balance Monthly P1 | Readout 50 G Weight Q6.P1 | Standard Acceptance Criteria Readout 50 G Weight Q6.P1 Between 49.99 to 50.01
IDT Electronic Balance Monthly | IDT Electronic Balance Monthly P1 | Readout 200 G Weight Q7.P1 | Standard Acceptance Criteria Readout 200 G Weight Q7.P1 Between 199.9 to 200.1
IDT Electronic Balance Monthly | IDT Electronic Balance Monthly P1 | Problem with Equipment Q8.P1 | Standard Acceptance Criteria Problem with Equipment Q8.P1 Equal to No
CM/Fixed Sites Electronic Balance Monthly | Electronic Balance DO Monthly P4 | Readout (g) After Tared Q1.P4 | Standard Acceptance Criteria Readout (g) After Tared Q1.P4 Between -0.01 to 0.01
CM/Fixed Sites Electronic Balance Monthly | Electronic Balance DO Monthly P4 | Readout 300 G Weight Q2.P4 | Standard Acceptance Criteria Readout 300 G Weight Q2.P4 Between 299.9 to 300.1
CM/Fixed Sites Electronic Balance Monthly | Electronic Balance DO Monthly P4 | Readout 500 G Weight Q3.P4 | Standard Acceptance Criteria Readout 500 G Weight Q3.P4 Between 499.9 to 500.1
CM/Fixed Sites Electronic Balance Monthly | Electronic Balance DO Monthly P4 | Readout 1000 G Weight Q4.P4 | Standard Acceptance Criteria Readout 1000 G Weight Q4.P4 Between 999.9 to 1000.1
CM/Fixed Sites Electronic Balance Monthly | Electronic Balance DO Monthly P4 | Problem with Equipment Q5.P4 | Standard Acceptance Criteria Problem with Equipment Q5.P4 Equal to No
Sorvall RC3BP+ Weekly | QC Performed P1 | Was the QC Performed Q1.P1 | Standard Acceptance Criteria Was the QC Performed Q1.P1 Equal to Yes
Sorvall RC3BP+ Weekly | Sorvall RC3BP+ Weekly P1.P2(Con) | Wash Rotor Chamber Q1.P1.P2(Con) | Standard Acceptance Criteria Wash Rotor Chamber Q1.P1.P2(Con) Equal to Performed
Sorvall RC3BP+ Weekly | Sorvall RC3BP+ Weekly P1.P2(Con) | Clean Cabinet Panels Q2.P1.P2(Con) | Standard Acceptance Criteria Clean Cabinet Panels Q2.P1.P2(Con) Equal to Performed
Sorvall RC3BP+ Weekly | Sorvall RC3BP+ Weekly P1.P2(Con) | Clean the Top Deck and Chamber Door Q3.P1.P2(Con) | Standard Acceptance Criteria Clean the Top Deck and Chamber Door Q3.P1.P2(Con) Equal to Performed
Sorvall RC3BP+ Weekly | Sorvall RC3BP+ Weekly P1.P2(Con) | Lubricate Cups Q4.P1.P2(Con) | Standard Acceptance Criteria Lubricate Cups Q4.P1.P2(Con) Equal to Performed
Sorvall RC3BP+ Weekly | Sorvall RC3BP+ Weekly P1.P2(Con) | Problem with Equipment Q5.P1.P2(Con) | Standard Acceptance Criteria Problem with Equipment Q5.P1.P2(Con) Equal to No
Manual Temperatures (Take Every 4 Hours) | Manual Temperature - Location P3.P2(Con) | Record Room Temperature - Reagent, Storage, Blood Component Q2.P3.P2(Con) | Standard Acceptance Criteria Record Room Temperature - Reagent, Storage, Blood Component Q2.P3.P2(Con) Between 20 to 24
Manual Temperatures (Take Every 4 Hours) | Manual Temperature - Location P3.P2(Con) | Problem with Equipment Q3.P3.P2(Con) | Standard Acceptance Criteria Problem with Equipment Q3.P3.P2(Con) Equal to No
Manual Temperatures (Take Every 4 Hours) | Manual Temperature Equipment P3.P3(Con) | Problem with Equipment Q3.P3.P3(Con) | Standard Acceptance Criteria Problem with Equipment Q3.P3.P3(Con) Equal to No
Sorvall RC3BP+ As Needed | Sorvall RC3BP+ As needed P1 | Disinfect the Rotor Q1.P1 | Standard Acceptance Criteria Disinfect the Rotor Q1.P1 Equal to Performed
Sorvall RC3BP+ As Needed | Sorvall RC3BP+ As needed P1 | Defrost the Rotor Chamber Q2.P1 | Standard Acceptance Criteria Defrost the Rotor Chamber Q2.P1 Equal to Performed
Sorvall RC3BP+ As Needed | Sorvall RC3BP+ As needed P1 | Wash the Spindle Q3.P1 | Standard Acceptance Criteria Wash the Spindle Q3.P1 Equal to Performed
Sorvall RC3BP+ As Needed | Sorvall RC3BP+ As needed P1 | Check the Gas Springs Q4.P1 | Standard Acceptance Criteria Check the Gas Springs Q4.P1 Equal to Performed
Sorvall RC3BP+ As Needed | Sorvall RC3BP+ As needed P1 | Check the Gas Spring Bracket Q5.P1 | Standard Acceptance Criteria Check the Gas Spring Bracket Q5.P1 Equal to Performed
Sorvall RC3BP+ As Needed | Sorvall RC3BP+ As needed P1 | Problem with Equipment Q6.P1 | Standard Acceptance Criteria Problem with Equipment Q6.P1 Equal to No
Sickledex Daily | QC Performed P1 | Was the QC Performed Q1.P1 | Standard Acceptance Criteria Was the QC Performed Q1.P1 Equal to No
Sickledex Daily | Sickledex Daily P1.P2(Con) | Positive Control Open Bottle Expiration Date Q1.P1.P2(Con) | Standard Acceptance Criteria Positive Control Open Bottle Expiration Date Q1.P1.P2(Con) Greater than or equal to Today
Sickledex Daily | Sickledex Daily P1.P2(Con) | Negative Control Open Bottle Expiration Date Q2.P1.P2(Con) | Standard Acceptance Criteria Negative Control Open Bottle Expiration Date Q2.P1.P2(Con) Greater than or equal to Today
Sickledex Daily | Sickledex Daily P1.P2(Con) | Sickledex Solution Open Bottle Expiration Date Q3.P1.P2(Con) | Standard Acceptance Criteria Sickledex Solution Open Bottle Expiration Date Q3.P1.P2(Con) Greater than or equal to Today
Sickledex Daily | Sickledex Daily P1.P2(Con) | Sickle Positive Control Result Q4.P1.P2(Con) | Standard Acceptance Criteria Sickle Positive Control Result Q4.P1.P2(Con) Equal to Positive - Acceptable
Sickledex Daily | Sickledex Daily P1.P2(Con) | Sickle Negative Control Result Q5.P1.P2(Con) | Standard Acceptance Criteria Sickle Negative Control Result Q5.P1.P2(Con) Equal to Negative - Acceptable
Sickledex Daily | Sickledex Daily P1.P2(Con) | Problem with QC Q6.P1.P2(Con) | Standard Acceptance Criteria Problem with QC Q6.P1.P2(Con) Equal to No
Trima Daily | Trima Daily P1 | Start-Up Q1.P1 | Standard Acceptance Criteria Start-Up Q1.P1 Equal to Performed - OK
Trima Daily | Trima Daily P1 | Problem with Equipment Q2.P1 | Standard Acceptance Criteria Problem with Equipment Q2.P1 Equal to No
Platelet Monthly | Problem with QC P1 | Problem with QC Q1.P1 | Standard Acceptance Criteria Problem with QC Q1.P1 Equal to No
Vital Sign Monitor/Manual BP Cuffs Daily | Vital Sign Monitor Daily P1 | Start-Up Check Q1.P1 | Standard Acceptance Criteria Start-Up Check Q1.P1 Equal to Performed - OK
Vital Sign Monitor/Manual BP Cuffs Daily | Vital Sign Monitor Daily P1 | Problem with Equipment Q2.P1 | Standard Acceptance Criteria Problem with Equipment Q2.P1 Equal to No
Vital Sign Monitor/Manual BP Cuffs Daily | Manual BP Cuffs Daily P2 | Acceptable for Use Q2.P2 | Standard Acceptance Criteria Acceptable for Use Q2.P2 Equal to Yes
Millipore Daily | Millipore Type I Water Daily P1 | Recirculated Water Q1.P1 | Standard Acceptance Criteria Recirculated Water Q1.P1 Equal to Performed - OK
Millipore Daily | Millipore Type I Water Daily P1 | Problem with Equipment Q2.P1 | Standard Acceptance Criteria Problem with Equipment Q2.P1 Equal to No
Equipment Problem Log (backup) | 11-Re-activate Out of Service Equipment P1.P2(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P4(Con) | Re-activate Out of Service Equipment Q1.P1.P2(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P4(Con) | Standard Acceptance Criteria Re-activate Out of Service Equipment Q1.P1.P2(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P4(Con) Equal to Instructions Acknowledged
Equipment Problem Log (backup) | 11-Re-activate Out of Service Equipment P1.P2(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P4(Con) | Re-activate Out of Service Equipment Q1.P1.P2(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P4(Con) | Standard Acceptance Criteria Re-activate Out of Service Equipment Q1.P1.P2(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P4(Con) Equal to Instructions Acknowledged
Equipment Problem Log (backup) | 11-Re-activate Out of Service Equipment P1.P2(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P4(Con) | Re-activate Out of Service Equipment Q1.P1.P2(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P4(Con) | Standard Acceptance Criteria Re-activate Out of Service Equipment Q1.P1.P2(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P4(Con) Equal to Instructions Acknowledged
Equipment Problem Log (backup) | 11-Re-activate Out of Service Equipment P1.P2(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P2(Con).P3(Con).P2(Con).P4(Con) | Re-activate Out of Service Equipment Q1.P1.P2(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P2(Con).P3(Con).P2(Con).P4(Con) | Standard Acceptance Criteria Re-activate Out of Service Equipment Q1.P1.P2(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P2(Con).P3(Con).P2(Con).P4(Con) Equal to Instructions Acknowledged
Equipment Problem Log (backup) | 11-Re-activate Out of Service Equipment P1.P2(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P2(Con).P4(Con) | Re-activate Out of Service Equipment Q1.P1.P2(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P2(Con).P4(Con) | Standard Acceptance Criteria Re-activate Out of Service Equipment Q1.P1.P2(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P2(Con).P4(Con) Equal to Instructions Acknowledged
Equipment Problem Log (backup) | 11-Re-activate Out of Service Equipment P1.P2(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P4(Con) | Re-activate Out of Service Equipment Q1.P1.P2(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P4(Con) | Standard Acceptance Criteria Re-activate Out of Service Equipment Q1.P1.P2(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P4(Con) Equal to Instructions Acknowledged
Equipment Problem Log (backup) | 11-Re-activate Out of Service Equipment P1.P2(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P4(Con) | Re-activate Out of Service Equipment Q1.P1.P2(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P4(Con) | Standard Acceptance Criteria Re-activate Out of Service Equipment Q1.P1.P2(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P4(Con) Equal to Instructions Acknowledged
Equipment Problem Log (backup) | 11-Re-activate Out of Service Equipment P1.P2(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P4(Con) | Re-activate Out of Service Equipment Q1.P1.P2(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P4(Con) | Standard Acceptance Criteria Re-activate Out of Service Equipment Q1.P1.P2(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P4(Con) Equal to Instructions Acknowledged
Equipment Problem Log (backup) | 11-Re-activate Out of Service Equipment P1.P2(Con).P2(Con).P3(Con).P2(Con).P4(Con) | Re-activate Out of Service Equipment Q1.P1.P2(Con).P2(Con).P3(Con).P2(Con).P4(Con) | Standard Acceptance Criteria Re-activate Out of Service Equipment Q1.P1.P2(Con).P2(Con).P3(Con).P2(Con).P4(Con) Equal to Instructions Acknowledged
Equipment Problem Log (backup) | 11-Re-activate Out of Service Equipment P1.P2(Con).P2(Con).P4(Con) | Re-activate Out of Service Equipment Q1.P1.P2(Con).P2(Con).P4(Con) | Standard Acceptance Criteria Re-activate Out of Service Equipment Q1.P1.P2(Con).P2(Con).P4(Con) Equal to Instructions Acknowledged
Equipment Problem Log (backup) | 11-Re-activate Out of Service Equipment P1.P2(Con).P3(Con).P2(Con).P2(Con).P2(Con).P2(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P4(Con) | Re-activate Out of Service Equipment Q1.P1.P2(Con).P3(Con).P2(Con).P2(Con).P2(Con).P2(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P4(Con) | Standard Acceptance Criteria Re-activate Out of Service Equipment Q1.P1.P2(Con).P3(Con).P2(Con).P2(Con).P2(Con).P2(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P4(Con) Equal to Instructions Acknowledged
Equipment Problem Log (backup) | 11-Re-activate Out of Service Equipment P1.P2(Con).P3(Con).P2(Con).P2(Con).P2(Con).P2(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P4(Con) | Re-activate Out of Service Equipment Q1.P1.P2(Con).P3(Con).P2(Con).P2(Con).P2(Con).P2(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P4(Con) | Standard Acceptance Criteria Re-activate Out of Service Equipment Q1.P1.P2(Con).P3(Con).P2(Con).P2(Con).P2(Con).P2(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P4(Con) Equal to Instructions Acknowledged
Equipment Problem Log (backup) | 11-Re-activate Out of Service Equipment P1.P2(Con).P3(Con).P2(Con).P2(Con).P2(Con).P2(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P4(Con) | Re-activate Out of Service Equipment Q1.P1.P2(Con).P3(Con).P2(Con).P2(Con).P2(Con).P2(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P4(Con) | Standard Acceptance Criteria Re-activate Out of Service Equipment Q1.P1.P2(Con).P3(Con).P2(Con).P2(Con).P2(Con).P2(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P4(Con) Equal to Instructions Acknowledged
Equipment Problem Log (backup) | 11-Re-activate Out of Service Equipment P1.P2(Con).P3(Con).P2(Con).P2(Con).P2(Con).P2(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P4(Con) | Re-activate Out of Service Equipment Q1.P1.P2(Con).P3(Con).P2(Con).P2(Con).P2(Con).P2(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P4(Con) | Standard Acceptance Criteria Re-activate Out of Service Equipment Q1.P1.P2(Con).P3(Con).P2(Con).P2(Con).P2(Con).P2(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P4(Con) Equal to Instructions Acknowledged
Equipment Problem Log (backup) | 11-Re-activate Out of Service Equipment P1.P2(Con).P3(Con).P2(Con).P2(Con).P2(Con).P2(Con).P2(Con).P3(Con).P2(Con).P4(Con) | Re-activate Out of Service Equipment Q1.P1.P2(Con).P3(Con).P2(Con).P2(Con).P2(Con).P2(Con).P2(Con).P3(Con).P2(Con).P4(Con) | Standard Acceptance Criteria Re-activate Out of Service Equipment Q1.P1.P2(Con).P3(Con).P2(Con).P2(Con).P2(Con).P2(Con).P2(Con).P3(Con).P2(Con).P4(Con) Equal to Instructions Acknowledged
Equipment Problem Log (backup) | 11-Re-activate Out of Service Equipment P1.P2(Con).P3(Con).P2(Con).P2(Con).P2(Con).P2(Con).P2(Con).P4(Con) | Re-activate Out of Service Equipment Q1.P1.P2(Con).P3(Con).P2(Con).P2(Con).P2(Con).P2(Con).P2(Con).P4(Con) | Standard Acceptance Criteria Re-activate Out of Service Equipment Q1.P1.P2(Con).P3(Con).P2(Con).P2(Con).P2(Con).P2(Con).P2(Con).P4(Con) Equal to Instructions Acknowledged
Equipment Problem Log (backup) | 11-Re-activate Out of Service Equipment P1.P2(Con).P3(Con).P2(Con).P2(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con) | Re-activate Out of Service Equipment Q1.P1.P2(Con).P3(Con).P2(Con).P2(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con) | Standard Acceptance Criteria Re-activate Out of Service Equipment Q1.P1.P2(Con).P3(Con).P2(Con).P2(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con) Equal to Instructions Acknowledged
Equipment Problem Log (backup) | 11-Re-activate Out of Service Equipment P1.P2(Con).P3(Con).P2(Con).P2(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P4(Con) | Re-activate Out of Service Equipment Q1.P1.P2(Con).P3(Con).P2(Con).P2(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P4(Con) | Standard Acceptance Criteria Re-activate Out of Service Equipment Q1.P1.P2(Con).P3(Con).P2(Con).P2(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P4(Con) Equal to Instructions Acknowledged
Equipment Problem Log (backup) | 11-Re-activate Out of Service Equipment P1.P2(Con).P3(Con).P2(Con).P2(Con).P2(Con).P3(Con).P2(Con).P4(Con) | Re-activate Out of Service Equipment Q1.P1.P2(Con).P3(Con).P2(Con).P2(Con).P2(Con).P3(Con).P2(Con).P4(Con) | Standard Acceptance Criteria Re-activate Out of Service Equipment Q1.P1.P2(Con).P3(Con).P2(Con).P2(Con).P2(Con).P3(Con).P2(Con).P4(Con) Equal to Instructions Acknowledged
Equipment Problem Log (backup) | 11-Re-activate Out of Service Equipment P1.P2(Con).P3(Con).P2(Con).P2(Con).P3(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P4(Con) | Re-activate Out of Service Equipment Q1.P1.P2(Con).P3(Con).P2(Con).P2(Con).P3(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P4(Con) | Standard Acceptance Criteria Re-activate Out of Service Equipment Q1.P1.P2(Con).P3(Con).P2(Con).P2(Con).P3(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P4(Con) Equal to Instructions Acknowledged
Equipment Problem Log (backup) | 11-Re-activate Out of Service Equipment P1.P2(Con).P3(Con).P2(Con).P2(Con).P3(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P4(Con) | Re-activate Out of Service Equipment Q1.P1.P2(Con).P3(Con).P2(Con).P2(Con).P3(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P4(Con) | Standard Acceptance Criteria Re-activate Out of Service Equipment Q1.P1.P2(Con).P3(Con).P2(Con).P2(Con).P3(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P4(Con) Equal to Instructions Acknowledged
Equipment Problem Log (backup) | 11-Re-activate Out of Service Equipment P1.P2(Con).P3(Con).P2(Con).P2(Con).P3(Con).P3(Con).P2(Con).P3(Con).P2(Con).P4(Con) | Re-activate Out of Service Equipment Q1.P1.P2(Con).P3(Con).P2(Con).P2(Con).P3(Con).P3(Con).P2(Con).P3(Con).P2(Con).P4(Con) | Standard Acceptance Criteria Re-activate Out of Service Equipment Q1.P1.P2(Con).P3(Con).P2(Con).P2(Con).P3(Con).P3(Con).P2(Con).P3(Con).P2(Con).P4(Con) Equal to Instructions Acknowledged
Equipment Problem Log (backup) | 11-Re-activate Out of Service Equipment P1.P2(Con).P3(Con).P2(Con).P2(Con).P3(Con).P3(Con).P2(Con).P4(Con) | Re-activate Out of Service Equipment Q1.P1.P2(Con).P3(Con).P2(Con).P2(Con).P3(Con).P3(Con).P2(Con).P4(Con) | Standard Acceptance Criteria Re-activate Out of Service Equipment Q1.P1.P2(Con).P3(Con).P2(Con).P2(Con).P3(Con).P3(Con).P2(Con).P4(Con) Equal to Instructions Acknowledged
Equipment Problem Log (backup) | 11-Re-activate Out of Service Equipment P1.P2(Con).P3(Con).P2(Con).P2(Con).P3(Con).P4(Con) | Re-activate Out of Service Equipment Q1.P1.P2(Con).P3(Con).P2(Con).P2(Con).P3(Con).P4(Con) | Standard Acceptance Criteria Re-activate Out of Service Equipment Q1.P1.P2(Con).P3(Con).P2(Con).P2(Con).P3(Con).P4(Con) Equal to Instructions Acknowledged
Equipment Problem Log (backup) | 11-Re-activate Out of Service Equipment P1.P2(Con).P3(Con).P3(Con).P2(Con).P2(Con).P2(Con).P4(Con) | Re-activate Out of Service Equipment Q1.P1.P2(Con).P3(Con).P3(Con).P2(Con).P2(Con).P2(Con).P4(Con) | Standard Acceptance Criteria Re-activate Out of Service Equipment Q1.P1.P2(Con).P3(Con).P3(Con).P2(Con).P2(Con).P2(Con).P4(Con) Equal to Instructions Acknowledged
Equipment Problem Log (backup) | 11-Re-activate Out of Service Equipment P1.P2(Con).P3(Con).P3(Con).P3(Con).P2(Con) | Re-activate Out of Service Equipment Q1.P1.P2(Con).P3(Con).P3(Con).P3(Con).P2(Con) | Standard Acceptance Criteria Re-activate Out of Service Equipment Q1.P1.P2(Con).P3(Con).P3(Con).P3(Con).P2(Con) Equal to Instructions Acknowledged
Equipment Problem Log (backup) | 11-Re-activate Out of Service Equipment P1.P3(Con).P2(Con).P2(Con).P2(Con).P2(Con).P2(Con) | Re-activate Out of Service Equipment Q1.P1.P3(Con).P2(Con).P2(Con).P2(Con).P2(Con).P2(Con) | Standard Acceptance Criteria Re-activate Out of Service Equipment Q1.P1.P3(Con).P2(Con).P2(Con).P2(Con).P2(Con).P2(Con) Equal to Instructions Acknowledged
Equipment Problem Log (backup) | 11-Re-activate Out of Service Equipment P1.P3(Con).P2(Con).P2(Con).P2(Con).P3(Con).P2(Con).P4(Con) | Re-activate Out of Service Equipment Q1.P1.P3(Con).P2(Con).P2(Con).P2(Con).P3(Con).P2(Con).P4(Con) | Standard Acceptance Criteria Re-activate Out of Service Equipment Q1.P1.P3(Con).P2(Con).P2(Con).P2(Con).P3(Con).P2(Con).P4(Con) Equal to Instructions Acknowledged
Equipment Problem Log (backup) | 11-Re-activate Out of Service Equipment P1.P3(Con).P2(Con).P2(Con).P3(Con).P4(Con) | Re-activate Out of Service Equipment Q1.P1.P3(Con).P2(Con).P2(Con).P3(Con).P4(Con) | Standard Acceptance Criteria Re-activate Out of Service Equipment Q1.P1.P3(Con).P2(Con).P2(Con).P3(Con).P4(Con) Equal to Instructions Acknowledged
Equipment Problem Log (backup) | 11-Re-activate Out of Service Equipment P1.P3(Con).P3(Con).P2(Con).P2(Con).P2(Con) | Re-activate Out of Service Equipment Q1.P1.P3(Con).P3(Con).P2(Con).P2(Con).P2(Con) | Standard Acceptance Criteria Re-activate Out of Service Equipment Q1.P1.P3(Con).P3(Con).P2(Con).P2(Con).P2(Con) Equal to Instructions Acknowledged
Equipment Problem Log (backup) | 11-Re-activate Out of Service Equipment P1.P3(Con).P3(Con).P2(Con).P3(Con).P2(Con).P2(Con) | Re-activate Out of Service Equipment Q1.P1.P3(Con).P3(Con).P2(Con).P3(Con).P2(Con).P2(Con) | Standard Acceptance Criteria Re-activate Out of Service Equipment Q1.P1.P3(Con).P3(Con).P2(Con).P3(Con).P2(Con).P2(Con) Equal to Instructions Acknowledged
Equipment Problem Log (backup) | 11-Re-activate Out of Service Equipment P1.P3(Con).P3(Con).P3(Con).P2(Con) | Re-activate Out of Service Equipment Q1.P1.P3(Con).P3(Con).P3(Con).P2(Con) | Standard Acceptance Criteria Re-activate Out of Service Equipment Q1.P1.P3(Con).P3(Con).P3(Con).P2(Con) Equal to Instructions Acknowledged
Equipment Problem Log (backup) | 11-Re-activate Out of Service Equipment P1.P3(Con).P3(Con).P3(Con).P3(Con).P2(Con).P2(Con) | Re-activate Out of Service Equipment Q1.P1.P3(Con).P3(Con).P3(Con).P3(Con).P2(Con).P2(Con) | Standard Acceptance Criteria Re-activate Out of Service Equipment Q1.P1.P3(Con).P3(Con).P3(Con).P3(Con).P2(Con).P2(Con) Equal to Instructions Acknowledged
Equipment Problem Log (backup) | 11-Re-activate Out of Service Equipment P1.P3(Con).P3(Con).P3(Con).P3(Con).P2(Con).P3(Con).P2(Con).P2(Con) | Re-activate Out of Service Equipment Q1.P1.P3(Con).P3(Con).P3(Con).P3(Con).P2(Con).P3(Con).P2(Con).P2(Con) | Standard Acceptance Criteria Re-activate Out of Service Equipment Q1.P1.P3(Con).P3(Con).P3(Con).P3(Con).P2(Con).P3(Con).P2(Con).P2(Con) Equal to Instructions Acknowledged
BacT Daily | BacT Daily P1 | Discarded 24 Hour No Growth Bottles? Q1.P1 | Standard Acceptance Criteria Discarded 24 Hour No Growth Bottles? Q1.P1 Equal to Performed
BacT Daily | BacT Daily P1 | Problem with Equipment Q2.P1 | Standard Acceptance Criteria Problem with Equipment Q2.P1 Equal to No
Tigris Daily | Tigris QC Performed P1 | Tigris QC Performed Q1.P1 | Standard Acceptance Criteria Tigris QC Performed Q1.P1 Equal to No - Not Needed
Tigris Daily | Tigris Daily P1.P2(Con) | Decontaminate Work Areas/Handled Surfaces Q1.P1.P2(Con) | Standard Acceptance Criteria Decontaminate Work Areas/Handled Surfaces Q1.P1.P2(Con) Equal to Performed
Tigris Daily | Tigris Daily P1.P2(Con) | Inspect Drip Pan for Drips/Leaks Q2.P1.P2(Con) | Standard Acceptance Criteria Inspect Drip Pan for Drips/Leaks Q2.P1.P2(Con) Equal to Performed
Tigris Daily | Tigris Daily P1.P2(Con) | Perform Tigris Daily Maintenance Q3.P1.P2(Con) | Standard Acceptance Criteria Perform Tigris Daily Maintenance Q3.P1.P2(Con) Equal to Acceptable
Tigris Daily | Tigris Daily P1.P2(Con) | Record Vacuum Pressure Q4.P1.P2(Con) | Standard Acceptance Criteria Record Vacuum Pressure Q4.P1.P2(Con) Between 6.5 to 12
Tigris Daily | Tigris Daily P1.P2(Con) | Clean DiTi Cone Q5.P1.P2(Con) | Standard Acceptance Criteria Clean DiTi Cone Q5.P1.P2(Con) Equal to Performed
Tigris Daily | Tigris Daily P1.P2(Con) | Inspect TCR Door Spring Q6.P1.P2(Con) | Standard Acceptance Criteria Inspect TCR Door Spring Q6.P1.P2(Con) Equal to Performed
Tigris Daily | Tigris Daily P1.P2(Con) | Inspect Fluids and Empty Waste Q7.P1.P2(Con) | Standard Acceptance Criteria Inspect Fluids and Empty Waste Q7.P1.P2(Con) Equal to Performed
Tigris Daily | Tigris Daily P1.P2(Con) | Problem with Equipment Q8.P1.P2(Con) | Standard Acceptance Criteria Problem with Equipment Q8.P1.P2(Con) Equal to No
RPI/RPI 250 Daily | QC Performed P1 | Was the QC Performed Q1.P1 | Standard Acceptance Criteria Was the QC Performed Q1.P1 Equal to No
RPI/RPI 250 Daily | RPI/RPI 250 Daily P1.P2(Con) | Record Average Temperature Q1.P1.P2(Con) | Standard Acceptance Criteria Record Average Temperature Q1.P1.P2(Con) Between 30 to 34
RPI/RPI 250 Daily | RPI/RPI 250 Daily P1.P2(Con) | Record Max Temperature Q2.P1.P2(Con) | Standard Acceptance Criteria Record Max Temperature Q2.P1.P2(Con) Less than or equal to 35
RPI/RPI 250 Daily | RPI/RPI 250 Daily P1.P2(Con) | Problem with Equipment Q3.P1.P2(Con) | Standard Acceptance Criteria Problem with Equipment Q3.P1.P2(Con) Equal to No
Panther Daily | Panther Daily P1 | Decontaminate Work Areas/Handled Surfaces Q1.P1 | Standard Acceptance Criteria Decontaminate Work Areas/Handled Surfaces Q1.P1 Equal to Performed
Panther Daily | Panther Daily P1 | Inspect Outside the Panther and Floor for Drips and Leaks Q2.P1 | Standard Acceptance Criteria Inspect Outside the Panther and Floor for Drips and Leaks Q2.P1 Equal to Performed - OK
Panther Daily | Panther Daily P1 | Perform Mag Wash Clean Q3.P1 | Standard Acceptance Criteria Perform Mag Wash Clean Q3.P1 Equal to Performed - OK
Panther Daily | Panther Daily P1 | Clean Sample Racks Q4.P1 | Standard Acceptance Criteria Clean Sample Racks Q4.P1 Equal to Performed
Panther Daily | Panther Daily P1 | Clean Reagent Racks Q5.P1 | Standard Acceptance Criteria Clean Reagent Racks Q5.P1 Equal to Performed
Panther Daily | Panther Daily P1 | Problem with Equipment Q6.P1 | Standard Acceptance Criteria Problem with Equipment Q6.P1 Equal to No
Equipment Problem Log | 4-Service Call P1.P2(Con).P4(Con).P2(Con) | Problem Resolved? Q3.P1.P2(Con).P4(Con).P2(Con) | Standard Acceptance Criteria Problem Resolved? Q3.P1.P2(Con).P4(Con).P2(Con) Equal to Yes
Equipment Problem Log | 6-Qualifications/Validation P1.P2(Con).P4(Con).P2(Con).P2(Con).P2(Con) | Qualification/Validation Acceptable? Q3.P1.P2(Con).P4(Con).P2(Con).P2(Con).P2(Con) | Advance Acceptnace Criteria Problem Resolved? Q3.P1.P2(Con).P4(Con).P2(Con) Equal to Yes Standard Acceptance Criteria Qualification/Validation Acceptable? Q3.P1.P2(Con).P4(Con).P2(Con).P2(Con).P2(Con) Equal to Yes
Equipment Problem Log | 11-Re-activate Out of Service Equipment P1.P2(Con).P4(Con).P2(Con).P2(Con).P2(Con).P5(Con).P2(Con).P2(Con) | Re-activate Out of Service Equipment Q1.P1.P2(Con).P4(Con).P2(Con).P2(Con).P2(Con).P5(Con).P2(Con).P2(Con) | Standard Acceptance Criteria Re-activate Out of Service Equipment Q1.P1.P2(Con).P4(Con).P2(Con).P2(Con).P2(Con).P5(Con).P2(Con).P2(Con) Equal to Instructions Acknowledged
Equipment Problem Log | 11-Re-activate Out of Service Equipment P1.P2(Con).P4(Con).P2(Con).P2(Con).P2(Con).P5(Con).P2(Con).P3(Con).P2(Con).P2(Con) | Re-activate Out of Service Equipment Q1.P1.P2(Con).P4(Con).P2(Con).P2(Con).P2(Con).P5(Con).P2(Con).P3(Con).P2(Con).P2(Con) | Standard Acceptance Criteria Re-activate Out of Service Equipment Q1.P1.P2(Con).P4(Con).P2(Con).P2(Con).P2(Con).P5(Con).P2(Con).P3(Con).P2(Con).P2(Con) Equal to Instructions Acknowledged
Equipment Problem Log | 11-Re-activate Out of Service Equipment P1.P2(Con).P4(Con).P2(Con).P2(Con).P2(Con).P5(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con) | Re-activate Out of Service Equipment Q1.P1.P2(Con).P4(Con).P2(Con).P2(Con).P2(Con).P5(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con) | Standard Acceptance Criteria Re-activate Out of Service Equipment Q1.P1.P2(Con).P4(Con).P2(Con).P2(Con).P2(Con).P5(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con) Equal to Instructions Acknowledged
Equipment Problem Log | 11-Re-activate Out of Service Equipment P1.P2(Con).P4(Con).P2(Con).P2(Con).P3(Con).P3(Con).P2(Con).P2(Con) | Re-activate Out of Service Equipment Q1.P1.P2(Con).P4(Con).P2(Con).P2(Con).P3(Con).P3(Con).P2(Con).P2(Con) | Standard Acceptance Criteria Re-activate Out of Service Equipment Q1.P1.P2(Con).P4(Con).P2(Con).P2(Con).P3(Con).P3(Con).P2(Con).P2(Con) Equal to Instructions Acknowledged
Equipment Problem Log | 4-Service Call P1.P3(Con).P2(Con) | Problem Resolved? Q3.P1.P3(Con).P2(Con) | Standard Acceptance Criteria Problem Resolved? Q3.P1.P3(Con).P2(Con) Equal to Yes
Equipment Problem Log | 6-Qualifications/Validation P1.P3(Con).P2(Con).P2(Con).P2(Con) | Qualification/Validation Acceptable? Q3.P1.P3(Con).P2(Con).P2(Con).P2(Con) | Standard Acceptance Criteria Qualification/Validation Acceptable? Q3.P1.P3(Con).P2(Con).P2(Con).P2(Con) Equal to Yes
Equipment Problem Log | 12-Supervisor Approval Re-qualification Final P1.P3(Con).P2(Con).P2(Con).P2(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con) | Approved by Supervisor Q2.P1.P3(Con).P2(Con).P2(Con).P2(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con) | Standard Acceptance Criteria Approved by Supervisor Q2.P1.P3(Con).P2(Con).P2(Con).P2(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con) Equal to Yes
Equipment Problem Log | 11-Re-activate Out of Service Equipment P1.P3(Con).P2(Con).P2(Con).P2(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P2(Con) | Re-activate Out of Service Equipment Q1.P1.P3(Con).P2(Con).P2(Con).P2(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P2(Con) | Standard Acceptance Criteria Re-activate Out of Service Equipment Q1.P1.P3(Con).P2(Con).P2(Con).P2(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P2(Con) Equal to Instructions Acknowledged
Equipment Problem Log | 11-Re-activate Out of Service Equipment P1.P3(Con).P2(Con).P2(Con).P2(Con).P3(Con).P2(Con).P2(Con) | Re-activate Out of Service Equipment Q1.P1.P3(Con).P2(Con).P2(Con).P2(Con).P3(Con).P2(Con).P2(Con) | Standard Acceptance Criteria Re-activate Out of Service Equipment Q1.P1.P3(Con).P2(Con).P2(Con).P2(Con).P3(Con).P2(Con).P2(Con) Equal to Instructions Acknowledged
Equipment Problem Log | 11-Re-activate Out of Service Equipment P1.P3(Con).P2(Con).P2(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P2(Con) | Re-activate Out of Service Equipment Q1.P1.P3(Con).P2(Con).P2(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P2(Con) | Standard Acceptance Criteria Re-activate Out of Service Equipment Q1.P1.P3(Con).P2(Con).P2(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P2(Con) Equal to Instructions Acknowledged
Equipment Problem Log | 11-Re-activate Out of Service Equipment P1.P3(Con).P2(Con).P2(Con).P3(Con).P3(Con).P2(Con).P2(Con) | Re-activate Out of Service Equipment Q1.P1.P3(Con).P2(Con).P2(Con).P3(Con).P3(Con).P2(Con).P2(Con) | Standard Acceptance Criteria Re-activate Out of Service Equipment Q1.P1.P3(Con).P2(Con).P2(Con).P3(Con).P3(Con).P2(Con).P2(Con) Equal to Instructions Acknowledged
Equipment Problem Log | 11-Re-activate Out of Service Equipment P1.P3(Con).P2(Con).P2(Con).P3(Con).P3(Con).P2(Con).P3(Con).P2(Con).P2(Con) | Re-activate Out of Service Equipment Q1.P1.P3(Con).P2(Con).P2(Con).P3(Con).P3(Con).P2(Con).P3(Con).P2(Con).P2(Con) | Standard Acceptance Criteria Re-activate Out of Service Equipment Q1.P1.P3(Con).P2(Con).P2(Con).P3(Con).P3(Con).P2(Con).P3(Con).P2(Con).P2(Con) Equal to Instructions Acknowledged
Equipment Problem Log | 11-Re-activate Out of Service Equipment P1.P3(Con).P2(Con).P2(Con).P3(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P2(Con) | Re-activate Out of Service Equipment Q1.P1.P3(Con).P2(Con).P2(Con).P3(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P2(Con) | Standard Acceptance Criteria Re-activate Out of Service Equipment Q1.P1.P3(Con).P2(Con).P2(Con).P3(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P2(Con) Equal to Instructions Acknowledged
Equipment Problem Log | 11-Re-activate Out of Service Equipment P1.P3(Con).P2(Con).P2(Con).P3(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P2(Con) | Re-activate Out of Service Equipment Q1.P1.P3(Con).P2(Con).P2(Con).P3(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P2(Con) | Standard Acceptance Criteria Re-activate Out of Service Equipment Q1.P1.P3(Con).P2(Con).P2(Con).P3(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P2(Con) Equal to Instructions Acknowledged
Equipment Problem Log | 11-Re-activate Out of Service Equipment P1.P3(Con).P2(Con).P2(Con).P3(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P2(Con) | Re-activate Out of Service Equipment Q1.P1.P3(Con).P2(Con).P2(Con).P3(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P2(Con) | Standard Acceptance Criteria Re-activate Out of Service Equipment Q1.P1.P3(Con).P2(Con).P2(Con).P3(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P2(Con) Equal to Instructions Acknowledged
Equipment Problem Log | 11-Re-activate Out of Service Equipment P1.P3(Con).P2(Con).P2(Con).P3(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P2(Con) | Re-activate Out of Service Equipment Q1.P1.P3(Con).P2(Con).P2(Con).P3(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P2(Con) | Standard Acceptance Criteria Re-activate Out of Service Equipment Q1.P1.P3(Con).P2(Con).P2(Con).P3(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P2(Con) Equal to Instructions Acknowledged
Equipment Problem Log | 11-Re-activate Out of Service Equipment P1.P3(Con).P2(Con).P2(Con).P3(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P2(Con) | Re-activate Out of Service Equipment Q1.P1.P3(Con).P2(Con).P2(Con).P3(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P2(Con) | Standard Acceptance Criteria Re-activate Out of Service Equipment Q1.P1.P3(Con).P2(Con).P2(Con).P3(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P2(Con) Equal to Instructions Acknowledged
Equipment Problem Log | 11-Re-activate Out of Service Equipment P1.P3(Con).P2(Con).P2(Con).P3(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con) | Re-activate Out of Service Equipment Q1.P1.P3(Con).P2(Con).P2(Con).P3(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con) | Standard Acceptance Criteria Re-activate Out of Service Equipment Q1.P1.P3(Con).P2(Con).P2(Con).P3(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con) Equal to Instructions Acknowledged
IDT Electronic Balance Daily | IDT Electronic Balance Daily QC P2 | Weight in Grams Q2.P2 | Standard Acceptance Criteria Weight in Grams Q2.P2 Between 199.9 to 200.1
IDT Electronic Balance Daily | IDT Electronic Balance Daily QC P2 | Problem with Equipment Q1.P2 | Standard Acceptance Criteria Problem with Equipment Q1.P2 Equal to No
Tigris Weekly | QC Performed P1 | Was the QC Performed Q1.P1 | Standard Acceptance Criteria Was the QC Performed Q1.P1 Equal to No
Tigris Weekly | Tigris Weekly P1.P2(Con) | Perform Advanced Magnetic Wash Aspirator Cleaning Q1.P1.P2(Con) | Standard Acceptance Criteria Perform Advanced Magnetic Wash Aspirator Cleaning Q1.P1.P2(Con) Equal to Performed - OK
Tigris Weekly | Tigris Weekly P1.P2(Con) | Shut Down and Restart System Q2.P1.P2(Con) | Standard Acceptance Criteria Shut Down and Restart System Q2.P1.P2(Con) Equal to Performed - OK
Tigris Weekly | Tigris Weekly P1.P2(Con) | Problem with Equipment Q3.P1.P2(Con) | Standard Acceptance Criteria Problem with Equipment Q3.P1.P2(Con) Equal to No
Electronic Balance Weekly | Electronic Balance Weekly P1 | Cleaned Electronic Scale Q1.P1 | Standard Acceptance Criteria Cleaned Electronic Scale Q1.P1 Equal to Performed - OK
Electronic Balance Weekly | Electronic Balance Weekly P1 | Problem with Equipment Q2.P1 | Standard Acceptance Criteria Problem with Equipment Q2.P1 Equal to No
Eyewash Weekly-inactive | Eyewash Weekly - Donor Operations P1 | Eyewash Unobstructed? Q1.P1 | Standard Acceptance Criteria Eyewash Unobstructed? Q1.P1 Equal to Yes
Eyewash Weekly-inactive | Eyewash Weekly - Donor Operations P1 | Eyewash Covers? Q2.P1 | Standard Acceptance Criteria Eyewash Covers? Q2.P1 Equal to Yes
Eyewash Weekly-inactive | Eyewash Weekly - Donor Operations P1 | Eyewash Functions/Activates? Q3.P1 | Standard Acceptance Criteria Eyewash Functions/Activates? Q3.P1 Equal to Yes
Eyewash Weekly-inactive | Eyewash Weekly - Donor Operations P1 | Eyewash Drainage? Q4.P1 | Standard Acceptance Criteria Eyewash Drainage? Q4.P1 Not equal to No
Eyewash Weekly-inactive | Eyewash Weekly - Donor Operations P1 | Problem with Equipment Q6.P1 | Standard Acceptance Criteria Problem with Equipment Q6.P1 Equal to No
Safety Shower Weekly | Safety Shower Weekly P1 | Activation Q1.P1 | Standard Acceptance Criteria Activation Q1.P1 Equal to Performed - OK
Safety Shower Weekly | Safety Shower Weekly P1 | Problem with Equipment Q2.P1 | Standard Acceptance Criteria Problem with Equipment Q2.P1 Equal to No
Panther Weekly | QC Performed P1 | Was the QC Performed Q1.P1 | Standard Acceptance Criteria Was the QC Performed Q1.P1 Equal to No
Panther Weekly | Panther Weekly P1.P2(Con) | Clean Sample Shield Q1.P1.P2(Con) | Standard Acceptance Criteria Clean Sample Shield Q1.P1.P2(Con) Equal to Performed
Panther Weekly | Panther Weekly P1.P2(Con) | Shutdown and Restart the System/PC Q2.P1.P2(Con) | Standard Acceptance Criteria Shutdown and Restart the System/PC Q2.P1.P2(Con) Equal to Performed - OK
Panther Weekly | Panther Weekly P1.P2(Con) | Problem with Equipment Q3.P1.P2(Con) | Standard Acceptance Criteria Problem with Equipment Q3.P1.P2(Con) Equal to No
duplicate of monthly - RPI Weekly | QC Performed P1 | Was the QC Performed Q1.P1 | Standard Acceptance Criteria Was the QC Performed Q1.P1 Equal to Yes
duplicate of monthly - RPI Weekly | RPI Weekly P1.P2(Con) | Clean RPI Q1.P1.P2(Con) | Standard Acceptance Criteria Clean RPI Q1.P1.P2(Con) Equal to Performed
duplicate of monthly - RPI Weekly | RPI Weekly P1.P2(Con) | Problem with Equipment Q2.P1.P2(Con) | Standard Acceptance Criteria Problem with Equipment Q2.P1.P2(Con) Equal to No
Millipore Monthly | Millipore Type I Water Monthly P1 | Resistivity from Q-Pod Arm @ 25˚C Q3.P1 | Standard Acceptance Criteria Resistivity from Q-Pod Arm @ 25˚C Q3.P1 Greater than 18
Millipore Monthly | Millipore Type I Water Monthly P1 | 5 Day CFU Count From Micro Bacterial Growth Q4.P1 | Standard Acceptance Criteria 5 Day CFU Count From Micro Bacterial Growth Q4.P1 Less than 1
Millipore Monthly | Millipore Type I Water Monthly P1 | Problem with Equipment Q5.P1 | Standard Acceptance Criteria Problem with Equipment Q5.P1 Equal to No
Tigris Monthly | Tigris Monthly P1 | Clean Workstation Q1.P1 | Standard Acceptance Criteria Clean Workstation Q1.P1 Equal to Performed
Tigris Monthly | Tigris Monthly P1 | Clean AR Bay Q2.P1 | Standard Acceptance Criteria Clean AR Bay Q2.P1 Equal to Performed - OK
Tigris Monthly | Tigris Monthly P1 | Clean Sample Bay Q3.P1 | Standard Acceptance Criteria Clean Sample Bay Q3.P1 Equal to Performed - OK
Tigris Monthly | Tigris Monthly P1 | Clean Fluid and Waste Drawers Q4.P1 | Standard Acceptance Criteria Clean Fluid and Waste Drawers Q4.P1 Equal to Performed
Tigris Monthly | Tigris Monthly P1 | Clean Waste Containers Q5.P1 | Standard Acceptance Criteria Clean Waste Containers Q5.P1 Equal to Performed
Tigris Monthly | Tigris Monthly P1 | Clean External Surfaces Q6.P1 | Standard Acceptance Criteria Clean External Surfaces Q6.P1 Equal to Performed
Tigris Monthly | Tigris Monthly P1 | Clean Pipettor Fluid Lines Q7.P1 | Standard Acceptance Criteria Clean Pipettor Fluid Lines Q7.P1 Equal to Performed - OK
Tigris Monthly | Tigris Monthly P1 | Clean System Fluid Bottles Q8.P1 | Standard Acceptance Criteria Clean System Fluid Bottles Q8.P1 Equal to Performed - OK
Tigris Monthly | Tigris Monthly P1 | Archive and Purge Worklists and Log Files Q9.P1 | Standard Acceptance Criteria Archive and Purge Worklists and Log Files Q9.P1 Equal to Performed - OK
Tigris Monthly | Tigris Monthly P1 | Problem with Equipment Q10.P1 | Standard Acceptance Criteria Problem with Equipment Q10.P1 Equal to No
RPI/RPI 250 Monthly | RPI/RPI 250 Monthly P1 | Clean RPI Q1.P1 | Standard Acceptance Criteria Clean RPI Q1.P1 Equal to Performed
RPI/RPI 250 Monthly | RPI/RPI 250 Monthly P1 | Problem with Equipment Q2.P1 | Standard Acceptance Criteria Problem with Equipment Q2.P1 Equal to No
Panther Monthly | Panther Monthly P1 | Clean Monitor and Scanner Q1.P1 | Standard Acceptance Criteria Clean Monitor and Scanner Q1.P1 Equal to Performed
Panther Monthly | Panther Monthly P1 | Clean Tip Drawer Q2.P1 | Standard Acceptance Criteria Clean Tip Drawer Q2.P1 Equal to Performed
Panther Monthly | Panther Monthly P1 | Clean Reagent Bay-T Q3.P1 | Standard Acceptance Criteria Clean Reagent Bay-T Q3.P1 Equal to Performed
Panther Monthly | Panther Monthly P1 | Clean Sample Bay-P Q8.P1 | Standard Acceptance Criteria Clean Sample Bay-P Q8.P1 Equal to Performed
Panther Monthly | Panther Monthly P1 | Clean Universal Fluids Drawer Q4.P1 | Standard Acceptance Criteria Clean Universal Fluids Drawer Q4.P1 Equal to Performed
Panther Monthly | Panther Monthly P1 | Clean Waste Drawer Q5.P1 | Standard Acceptance Criteria Clean Waste Drawer Q5.P1 Equal to Performed
Panther Monthly | Panther Monthly P1 | Clean External Surfaces Q6.P1 | Standard Acceptance Criteria Clean External Surfaces Q6.P1 Equal to Performed
Panther Monthly | Panther Monthly P1 | Backup Database Q7.P1 | Standard Acceptance Criteria Backup Database Q7.P1 Equal to Performed - OK
Panther Monthly | Panther Monthly P1 | Problem with Equipment Q9.P1 | Standard Acceptance Criteria Problem with Equipment Q9.P1 Equal to No
Tigris PM Bi-Annual | PM P1 | PM Performed Q3.P1 | Standard Acceptance Criteria PM Performed Q3.P1 Equal to Performed - OK
Tigris PM Bi-Annual | PM P1 | Problem with Equipment Q4.P1 | Standard Acceptance Criteria Problem with Equipment Q4.P1 Equal to No
Millipore Annual PM Vendor | PM P1 | Preventative Maintenance Performed Q3.P1 | Standard Acceptance Criteria Preventative Maintenance Performed Q3.P1 Equal to Performed - OK
Millipore Annual PM Vendor | PM P1 | Problem with Equipment Q4.P1 | Standard Acceptance Criteria Problem with Equipment Q4.P1 Equal to No
Millipore PM Annual | PM P1 | PM Performed Q3.P1 | Standard Acceptance Criteria PM Performed Q3.P1 Equal to Performed - OK
Millipore PM Annual | PM P1 | Problem with Equipment Q4.P1 | Standard Acceptance Criteria Problem with Equipment Q4.P1 Equal to No
BacT PM Annual | PM P1 | PM Performed Q3.P1 | Standard Acceptance Criteria PM Performed Q3.P1 Equal to Performed - OK
BacT PM Annual | PM P1 | Problem with Equipment Q4.P1 | Standard Acceptance Criteria Problem with Equipment Q4.P1 Equal to No
Refrigerator PM Annual | PM P1 | PM Performed Q3.P1 | Standard Acceptance Criteria PM Performed Q3.P1 Equal to Performed - OK
Refrigerator PM Annual | PM P1 | Problem with Equipment Q4.P1 | Standard Acceptance Criteria Problem with Equipment Q4.P1 Equal to No
Microscope Maintenance Annual | Microscope Maintenance Annual P1 | Lubricate Gears? Q1.P1 | Standard Acceptance Criteria Lubricate Gears? Q1.P1 Equal to Performed
Microscope Maintenance Annual | Microscope Maintenance Annual P1 | Cleaned Objectives? Q2.P1 | Standard Acceptance Criteria Cleaned Objectives? Q2.P1 Equal to Performed
Microscope Maintenance Annual | Microscope Maintenance Annual P1 | Problem with Equipment Q3.P1 | Standard Acceptance Criteria Problem with Equipment Q3.P1 Equal to No
Electronic Balance NAT Annual PM | PM P1 | Preventative Maintenance Performed Q3.P1 | Standard Acceptance Criteria Preventative Maintenance Performed Q3.P1 Equal to Performed - OK
Electronic Balance NAT Annual PM | PM P1 | Problem with Equipment Q4.P1 | Standard Acceptance Criteria Problem with Equipment Q4.P1 Equal to No
Tigris As Needed | Tigris As Needed P1 | Clean Reagent Pipettor Probes Manually Q1.P1 | Standard Acceptance Criteria Clean Reagent Pipettor Probes Manually Q1.P1 Equal to Performed
Tigris As Needed | Tigris As Needed P1 | Perform Advanced Clean Pipettor Fluid Lines Q2.P1 | Standard Acceptance Criteria Perform Advanced Clean Pipettor Fluid Lines Q2.P1 Equal to Performed - OK
Tigris As Needed | Tigris As Needed P1 | Clean Mag Wash Aspirators Q3.P1 | Standard Acceptance Criteria Clean Mag Wash Aspirators Q3.P1 Equal to Performed - OK
Tigris As Needed | Tigris As Needed P1 | Check Pipettor Q4.P1 | Standard Acceptance Criteria Check Pipettor Q4.P1 Equal to Performed - OK
Tigris As Needed | Tigris As Needed P1 | Check Luminometer Q5.P1 | Standard Acceptance Criteria Check Luminometer Q5.P1 Equal to Performed - OK
Tigris As Needed | Tigris As Needed P1 | Idle Maintenance Q6.P1 | Standard Acceptance Criteria Idle Maintenance Q6.P1 Equal to Performed - OK
Tigris As Needed | Tigris As Needed P1 | Maintenance Unload Q7.P1 | Standard Acceptance Criteria Maintenance Unload Q7.P1 Equal to Performed - OK
Tigris As Needed | Tigris As Needed P1 | Problem with Equipment Q8.P1 | Standard Acceptance Criteria Problem with Equipment Q8.P1 Equal to No
Neo Residual Volume As Needed | Neo Residual Volume P1 | Problem with Equipment Q4.P1 | Standard Acceptance Criteria Problem with Equipment Q4.P1 Equal to No
Amicus Weekly | Amicus Weekly QC P1 | Inlet Air Filter Cleaned Q1.P1 | Standard Acceptance Criteria Inlet Air Filter Cleaned Q1.P1 Equal to Performed - OK
Amicus Weekly | Amicus Weekly QC P1 | Interface Detector System Cleaned Q2.P1 | Standard Acceptance Criteria Interface Detector System Cleaned Q2.P1 Equal to Performed - OK
Amicus Weekly | Amicus Weekly QC P1 | Touch Screen Cleaned Q3.P1 | Standard Acceptance Criteria Touch Screen Cleaned Q3.P1 Equal to Performed - OK
Amicus Weekly | Amicus Weekly QC P1 | Door Window Cleaned Q4.P1 | Standard Acceptance Criteria Door Window Cleaned Q4.P1 Equal to Performed - OK
Amicus Weekly | Amicus Weekly QC P1 | Problem with Equipment Q5.P1 | Standard Acceptance Criteria Problem with Equipment Q5.P1 Equal to No
Spectra Optia Weekly | Spectra Optia Weekly QC P4 | Surface Cleaned Q1.P4 | Standard Acceptance Criteria Surface Cleaned Q1.P4 Equal to Performed - OK
Spectra Optia Weekly | Spectra Optia Weekly QC P4 | Sensors Cleaned Q2.P4 | Standard Acceptance Criteria Sensors Cleaned Q2.P4 Equal to Performed - OK
Spectra Optia Weekly | Spectra Optia Weekly QC P4 | Detectors Cleaned Q3.P4 | Standard Acceptance Criteria Detectors Cleaned Q3.P4 Equal to Performed - OK
Spectra Optia Weekly | Spectra Optia Weekly QC P4 | Valves Cleaned Q4.P4 | Standard Acceptance Criteria Valves Cleaned Q4.P4 Equal to Performed - OK
Spectra Optia Weekly | Spectra Optia Weekly QC P4 | Problem with Equipment Q5.P4 | Standard Acceptance Criteria Problem with Equipment Q5.P4 Equal to No
Microscope Maintenance As Needed | Microscope As Needed P1 | Problem with Equipment Q3.P1 | Standard Acceptance Criteria Problem with Equipment Q3.P1 Equal to No
Panther As Needed | Panther As Needed P1 | Replace Mag Wash Solution Q1.P1 | Standard Acceptance Criteria Replace Mag Wash Solution Q1.P1 Equal to Performed - OK
Panther As Needed | Panther As Needed P1 | Problem with Equipment Q2.P1 | Standard Acceptance Criteria Problem with Equipment Q2.P1 Equal to No
Neo PipTest As Needed | Neo PipTest P1 | Strip 1: Actual Volume Q3.P1 | Standard Acceptance Criteria Strip 1: Actual Volume Q3.P1 Between 0.18 to 0.22
Neo PipTest As Needed | Neo PipTest P1 | Strip 2: Actual Volume Q6.P1 | Standard Acceptance Criteria Strip 2: Actual Volume Q6.P1 Between 0.36 to 0.44
Neo PipTest As Needed | Neo PipTest P1 | Strip 3: Actual Volume Q9.P1 | Standard Acceptance Criteria Strip 3: Actual Volume Q9.P1 Between 0.54 to 0.66
Neo PipTest As Needed | Neo PipTest P1 | Strip 4: Actual Volume Q12.P1 | Standard Acceptance Criteria Strip 4: Actual Volume Q12.P1 Between 0.72 to 0.88
Neo PipTest As Needed | Neo PipTest P1 | Strip 5: Actual Volume Q15.P1 | Standard Acceptance Criteria Strip 5: Actual Volume Q15.P1 Between 0.36 to 0.44
Neo PipTest As Needed | Neo PipTest P1 | Strip 6: Actual Volume Q18.P1 | Standard Acceptance Criteria Strip 6: Actual Volume Q18.P1 Between 0.36 to 0.44
Neo PipTest As Needed | Neo PipTest P1 | Strip 7: Actual Volume Q21.P1 | Standard Acceptance Criteria Strip 7: Actual Volume Q21.P1 Between 0.36 to 0.44
Neo PipTest As Needed | Neo PipTest P1 | Strip 8: Actual Volume Q24.P1 | Standard Acceptance Criteria Strip 8: Actual Volume Q24.P1 Between 0.36 to 0.44
Neo PipTest As Needed | Neo PipTest P1 | Strip 9: Actual Volume Q27.P1 | Standard Acceptance Criteria Strip 9: Actual Volume Q27.P1 Between 0.75 to 0.92
Neo PipTest As Needed | Neo PipTest P1 | Strip 10: Actual Volume Q30.P1 | Standard Acceptance Criteria Strip 10: Actual Volume Q30.P1 Between 0.75 to 0.92
Neo PipTest As Needed | Neo PipTest P1 | Strip 11: Actual Volume Q33.P1 | Standard Acceptance Criteria Strip 11: Actual Volume Q33.P1 Between 0.75 to 0.92
Neo PipTest As Needed | Neo PipTest P1 | Strip 12: Actual Volume Q36.P1 | Standard Acceptance Criteria Strip 12: Actual Volume Q36.P1 Between 0.75 to 0.92
Neo PipTest As Needed | Neo PipTest P1 | Problem with Equipment Q38.P1 | Standard Acceptance Criteria Problem with Equipment Q38.P1 Equal to No
LT Day Shift Daily | Clean and Disinfect Daily Log P1 | Clean and Disinfect ALL Working Surface Areas Q1.P1 | Standard Acceptance Criteria Clean and Disinfect ALL Working Surface Areas Q1.P1 Equal to Performed - OK
LT Day Shift Daily | Clean and Disinfect Daily Log P1 | Disinfect Pooling Pole Q2.P1 | Standard Acceptance Criteria Disinfect Pooling Pole Q2.P1 Equal to Performed - OK
LT Day Shift Daily | Clean and Disinfect Daily Log P1 | Disinfect Hematron Q3.P1 | Standard Acceptance Criteria Disinfect Hematron Q3.P1 Equal to Performed - OK
LT Day Shift Daily | Clean and Disinfect Daily Log P1 | Disinfect SCD Q4.P1 | Standard Acceptance Criteria Disinfect SCD Q4.P1 Equal to Performed - OK
LT Day Shift Daily | Clean and Disinfect Daily Log P1 | Clean Tube Station Q5.P1 | Standard Acceptance Criteria Clean Tube Station Q5.P1 Equal to Performed - OK
LT Evening Shift Daily | Clean and Disinfect P1 | Clean and Disinfect ALL Working Surface Areas Q1.P1 | Standard Acceptance Criteria Clean and Disinfect ALL Working Surface Areas Q1.P1 Equal to Performed - OK
LT Evening Shift Daily | Clean and Disinfect P1 | Clean Tube Station Q2.P1 | Standard Acceptance Criteria Clean Tube Station Q2.P1 Equal to Performed - OK
LT Night Shift Daily | Clean and Disinfect P1 | Clean and Disinfect ALL Working Surface Areas Q1.P1 | Standard Acceptance Criteria Clean and Disinfect ALL Working Surface Areas Q1.P1 Equal to Performed - OK
LT Night Shift Daily | Clean and Disinfect P1 | Clean Tube Station Q2.P1 | Standard Acceptance Criteria Clean Tube Station Q2.P1 Equal to Performed - OK
Weight Scale Daily | Weight Scale Daily P1 | Scale Reading Q1.P1 | Standard Acceptance Criteria Scale Reading Q1.P1 Between 49.8 to 50.2
Weight Scale Daily | Weight Scale Daily P1 | Problem with Equipment Q2.P1 | Standard Acceptance Criteria Problem with Equipment Q2.P1 Equal to No
Refrigerated Centrifuge Daily - DO NOT USE | Refrigerated Centrifuge Weekly P1 | Wash Rotor Chamber Q1.P1 | Standard Acceptance Criteria Wash Rotor Chamber Q1.P1 Equal to Performed - OK
Refrigerated Centrifuge Daily - DO NOT USE | Refrigerated Centrifuge Weekly P1 | Clean Cabinet Panels Q2.P1 | Standard Acceptance Criteria Clean Cabinet Panels Q2.P1 Equal to Performed - OK
Refrigerated Centrifuge Daily - DO NOT USE | Refrigerated Centrifuge Weekly P1 | Clean the Top Deck and Chamber Door Q3.P1 | Standard Acceptance Criteria Clean the Top Deck and Chamber Door Q3.P1 Equal to Performed - OK
Refrigerated Centrifuge Daily - DO NOT USE | Refrigerated Centrifuge Weekly P1 | Lubricate Cups Q4.P1 | Standard Acceptance Criteria Lubricate Cups Q4.P1 Equal to Performed - OK
Refrigerated Centrifuge Daily | Refrigerated Centrifuge Daily P1 | QC Run Q1.P1 | Standard Acceptance Criteria QC Run Q1.P1 Equal to Performed - OK
Refrigerated Centrifuge Daily | Refrigerated Centrifuge Daily P1 | Problem with Equipment Q2.P1 | Standard Acceptance Criteria Problem with Equipment Q2.P1 Equal to No
Prism Cut Off Verification | Prism Cut Off Verification - Day 1 P1 | Negative Control Valid? Q3.P1 | Standard Acceptance Criteria Negative Control Valid? Q3.P1 Equal to Yes
Prism Cut Off Verification | Prism Cut Off Verification - Day 1 P1 | Positive Control Valid? Q4.P1 | Standard Acceptance Criteria Positive Control Valid? Q4.P1 Equal to Yes
Prism Cut Off Verification | Prism Cut Off Verification - Day 1 P1 | Abbott A0_ Positive Control Valid? Q5.P1 | Standard Acceptance Criteria Abbott A0_ Positive Control Valid? Q5.P1 Equal to Yes
Prism Cut Off Verification | Prism Cut Off Verification - Day 1 P1 | Non-Reactive Valid? Q10.P1 | Standard Acceptance Criteria Non-Reactive Valid? Q10.P1 Equal to Yes
Prism Cut Off Verification | Prism Cut Off Verification - Day 1 P1 | Reactive Valid? Q13.P1 | Standard Acceptance Criteria Reactive Valid? Q13.P1 Equal to Yes
Prism Cut Off Verification | Prism Cut Off Verification - Day 2 P2 | Day 2 Negative Control Valid? Q3.P2 | Standard Acceptance Criteria Day 2 Negative Control Valid? Q3.P2 Equal to Yes
Prism Cut Off Verification | Prism Cut Off Verification - Day 2 P2 | Day 2 Positive Control Valid? Q4.P2 | Standard Acceptance Criteria Day 2 Positive Control Valid? Q4.P2 Equal to Yes
Prism Cut Off Verification | Prism Cut Off Verification - Day 2 P2 | Day 2 Abbott A0_ Positive Control Valid? Q5.P2 | Standard Acceptance Criteria Day 2 Abbott A0_ Positive Control Valid? Q5.P2 Equal to Yes
Prism Cut Off Verification | Prism Cut Off Verification - Day 2 P2 | Day 2 Non-Reactive Valid? Q10.P2 | Standard Acceptance Criteria Day 2 Non-Reactive Valid? Q10.P2 Equal to Yes
Prism Cut Off Verification | Prism Cut Off Verification - Day 2 P2 | Day 2 Reactive Valid? Q13.P2 | Standard Acceptance Criteria Day 2 Reactive Valid? Q13.P2 Equal to Yes
Prism Cut Off Verification | Prism Cut Off Verification P3 | Is Non-Reactive S/CO on Day 1 Close to Non-Reactive S/CO on Day 2 (are Both Non-reactive?) Q1.P3 | Standard Acceptance Criteria Is Non-Reactive S/CO on Day 1 Close to Non-Reactive S/CO on Day 2 (are Both Non-reactive?) Q1.P3 Equal to Yes
Prism Cut Off Verification | Prism Cut Off Verification P3 | Is Reactive S/CO on Day 1 Close to Reactive S/CO on Day 2 (are Both Reactive?) Q2.P3 | Standard Acceptance Criteria Is Reactive S/CO on Day 1 Close to Reactive S/CO on Day 2 (are Both Reactive?) Q2.P3 Equal to Yes
Prism Cut Off Verification | Prism Cut Off Verification P3 | Are the Runs Acceptable? (All Controls/calibrators Results Are as Expected?) Q3.P3 | Standard Acceptance Criteria Are the Runs Acceptable? (All Controls/calibrators Results Are as Expected?) Q3.P3 Equal to Acceptable
Prism Cut Off Verification | Prism Cut Off Verification P3 | Problem with Equipment Q4.P3 | Standard Acceptance Criteria Problem with Equipment Q4.P3 Equal to No
Ultrio Cutoff/Correlation As Needed | Ultrio Cutoff/Correlation P1 | Ultrio Negative External Control Day 1 Valid Q4.P1 | Standard Acceptance Criteria Ultrio Negative External Control Day 1 Valid Q4.P1 Equal to Yes
Ultrio Cutoff/Correlation As Needed | Ultrio Cutoff/Correlation P1 | Ultrio HIV-1 Positive External Control Day 1 Valid Q5.P1 | Standard Acceptance Criteria Ultrio HIV-1 Positive External Control Day 1 Valid Q5.P1 Equal to Yes
Ultrio Cutoff/Correlation As Needed | Ultrio Cutoff/Correlation P1 | Ultrio HCV Positive External Control Day 1 Valid Q6.P1 | Standard Acceptance Criteria Ultrio HCV Positive External Control Day 1 Valid Q6.P1 Equal to Yes
Ultrio Cutoff/Correlation As Needed | Ultrio Cutoff/Correlation P1 | Ultrio HBV Positive External Control Day 1 Valid Q7.P1 | Standard Acceptance Criteria Ultrio HBV Positive External Control Day 1 Valid Q7.P1 Equal to Yes
Ultrio Cutoff/Correlation As Needed | Ultrio Cutoff/Correlation P1 | Non-Reactive Day 1 Valid Q11.P1 | Standard Acceptance Criteria Non-Reactive Day 1 Valid Q11.P1 Equal to Yes
Ultrio Cutoff/Correlation As Needed | Ultrio Cutoff/Correlation P1 | Reactive Day 1 Valid Q12.P1 | Standard Acceptance Criteria Reactive Day 1 Valid Q12.P1 Equal to Yes
Ultrio Cutoff/Correlation As Needed | Ultrio Cutoff/Correlation P1 | Ultrio Negative External Control Day 2 Valid Q13.P1 | Standard Acceptance Criteria Ultrio Negative External Control Day 2 Valid Q13.P1 Equal to Yes
Ultrio Cutoff/Correlation As Needed | Ultrio Cutoff/Correlation P1 | Ultrio HIV-1 Positive External Control Day 2 Valid Q14.P1 | Standard Acceptance Criteria Ultrio HIV-1 Positive External Control Day 2 Valid Q14.P1 Equal to Yes
Ultrio Cutoff/Correlation As Needed | Ultrio Cutoff/Correlation P1 | Ultrio HCV Positive External Control Day 2 Valid Q15.P1 | Standard Acceptance Criteria Ultrio HCV Positive External Control Day 2 Valid Q15.P1 Equal to Yes
Ultrio Cutoff/Correlation As Needed | Ultrio Cutoff/Correlation P1 | Ultrio HBV Positive External Control Day 2 Valid Q16.P1 | Standard Acceptance Criteria Ultrio HBV Positive External Control Day 2 Valid Q16.P1 Equal to Yes
Ultrio Cutoff/Correlation As Needed | Ultrio Cutoff/Correlation P1 | Non-Reactive Day 2 Valid Q20.P1 | Standard Acceptance Criteria Non-Reactive Day 2 Valid Q20.P1 Equal to Yes
Ultrio Cutoff/Correlation As Needed | Ultrio Cutoff/Correlation P1 | Reactive Day 2 Valid Q21.P1 | Standard Acceptance Criteria Reactive Day 2 Valid Q21.P1 Equal to Yes
Ultrio Cutoff/Correlation As Needed | Ultrio Cutoff/Correlation P1 | Is Non-Reactive Analyte RLU on Day 1 Close to Non-Reactive Analyte on Day 2 (are Both Non-reactive?) Q22.P1 | Standard Acceptance Criteria Is Non-Reactive Analyte RLU on Day 1 Close to Non-Reactive Analyte on Day 2 (are Both Non-reactive?) Q22.P1 Equal to Yes
Ultrio Cutoff/Correlation As Needed | Ultrio Cutoff/Correlation P1 | Is Reactive Analyte RLU on Day 1 Close to Reactive Analyte RLU on Day 2 (are Both Reactive?) Q23.P1 | Standard Acceptance Criteria Is Reactive Analyte RLU on Day 1 Close to Reactive Analyte RLU on Day 2 (are Both Reactive?) Q23.P1 Equal to Yes
Ultrio Cutoff/Correlation As Needed | Ultrio Cutoff/Correlation P1 | Are the Runs Acceptable? (All Controls/calibrators Results Are as Expected?) Q24.P1 | Standard Acceptance Criteria Are the Runs Acceptable? (All Controls/calibrators Results Are as Expected?) Q24.P1 Equal to Acceptable
Ultrio Cutoff/Correlation As Needed | Ultrio Cutoff/Correlation P1 | Is Correlation of Instruments Acceptable? (Answer Only If 2 Instruments Are Used) Q25.P1 | Standard Acceptance Criteria Is Correlation of Instruments Acceptable? (Answer Only If 2 Instruments Are Used) Q25.P1 Equal to Acceptable
Ultrio Cutoff/Correlation As Needed | Ultrio Cutoff/Correlation P1 | Problem with Equipment Q26.P1 | Standard Acceptance Criteria Problem with Equipment Q26.P1 Equal to No
Cell Washer Weekly | Cell Washer Weekly P1 | Decontamination Q1.P1 | Standard Acceptance Criteria Decontamination Q1.P1 Equal to Performed - OK
Waterbath Weekly | Waterbath Weekly P1 | Empty and Clean Q1.P1 | Standard Acceptance Criteria Empty and Clean Q1.P1 Equal to Performed - OK
Waterbath Monthly | Waterbath Monthly P1 | Lubricate Basket Assembly Q1.P1 | Standard Acceptance Criteria Lubricate Basket Assembly Q1.P1 Equal to Performed - OK
Waterbath Monthly | Waterbath Monthly P1 | Problem with Equipment Q2.P1 | Standard Acceptance Criteria Problem with Equipment Q2.P1 Equal to No
IBL 437 Irradiator Quarterly | IBL 437 Irradiator Quarterly P1 | Door Interlock Q1.P1 | Standard Acceptance Criteria Door Interlock Q1.P1 Equal to Passed
IBL 437 Irradiator Quarterly | IBL 437 Irradiator Quarterly P1 | Emergency Power Check Q2.P1 | Standard Acceptance Criteria Emergency Power Check Q2.P1 Equal to Passed
IBL 437 Irradiator Quarterly | IBL 437 Irradiator Quarterly P1 | Difference in Dosage Between Dosewriter and Chart Q5.P1 | Standard Acceptance Criteria Difference in Dosage Between Dosewriter and Chart Q5.P1 Between -0.01 to 0.01
IBL 437 Irradiator Quarterly | IBL 437 Irradiator Quarterly P1 | Timer Check Q6.P1 | Standard Acceptance Criteria Timer Check Q6.P1 Equal to Passed
IBL 437 Irradiator Quarterly | IBL 437 Irradiator Quarterly P1 | Problem with Equipment Q7.P1 | Standard Acceptance Criteria Problem with Equipment Q7.P1 Equal to No
Capture-P Washstation Weekly | Capture-P Washstation Weekly P1 | Backflush PBS and DI Reservoir Q1.P1 | Standard Acceptance Criteria Backflush PBS and DI Reservoir Q1.P1 Equal to Performed - OK
Capture-P Workstation Monthly | Capture-P Washstation Monthly P1 | Check Exhaust Filter Q1.P1 | Standard Acceptance Criteria Check Exhaust Filter Q1.P1 Equal to Performed - OK
Capture-P Workstation Monthly | Capture-P Washstation Monthly P1 | CSW 100 Dispense Volume Q2.P1 | Standard Acceptance Criteria CSW 100 Dispense Volume Q2.P1 Between 21.6 to 26.4
Capture-P Washstation Quarterly | Capture-P Workstation Quarterly P1 | Clean Manifold Needles Q1.P1 | Standard Acceptance Criteria Clean Manifold Needles Q1.P1 Equal to Performed - OK
Equipment Problem Log (copy) | 11-Re-activate Out of Service Equipment P1.P2(Con).P2(Con).P2(Con) | Re-activate Out of Service Equipment Q1.P1.P2(Con).P2(Con).P2(Con) | Standard Acceptance Criteria Re-activate Out of Service Equipment Q1.P1.P2(Con).P2(Con).P2(Con) Equal to Instructions Acknowledged
Equipment Problem Log (copy) | 11-Re-activate Out of Service Equipment P1.P2(Con).P2(Con).P3(Con).P2(Con).P2(Con) | Re-activate Out of Service Equipment Q1.P1.P2(Con).P2(Con).P3(Con).P2(Con).P2(Con) | Standard Acceptance Criteria Re-activate Out of Service Equipment Q1.P1.P2(Con).P2(Con).P3(Con).P2(Con).P2(Con) Equal to Instructions Acknowledged
Equipment Problem Log (copy) | 11-Re-activate Out of Service Equipment P1.P2(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P2(Con) | Re-activate Out of Service Equipment Q1.P1.P2(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P2(Con) | Standard Acceptance Criteria Re-activate Out of Service Equipment Q1.P1.P2(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P2(Con) Equal to Instructions Acknowledged
Equipment Problem Log (copy) | 4-Service Call P1.P2(Con).P4(Con).P2(Con) | Problem Resolved? Q3.P1.P2(Con).P4(Con).P2(Con) | Standard Acceptance Criteria Problem Resolved? Q3.P1.P2(Con).P4(Con).P2(Con) Equal to Yes
Equipment Problem Log (copy) | 11-Re-activate Out of Service Equipment P1.P2(Con).P4(Con).P2(Con).P2(Con).P2(Con).P4(Con).P2(Con) | Re-activate Out of Service Equipment Q1.P1.P2(Con).P4(Con).P2(Con).P2(Con).P2(Con).P4(Con).P2(Con) | Standard Acceptance Criteria Re-activate Out of Service Equipment Q1.P1.P2(Con).P4(Con).P2(Con).P2(Con).P2(Con).P4(Con).P2(Con) Equal to Instructions Acknowledged
Equipment Problem Log (copy) | 11-Re-activate Out of Service Equipment P1.P2(Con).P4(Con).P2(Con).P2(Con).P2(Con).P4(Con).P3(Con).P2(Con).P2(Con) | Re-activate Out of Service Equipment Q1.P1.P2(Con).P4(Con).P2(Con).P2(Con).P2(Con).P4(Con).P3(Con).P2(Con).P2(Con) | Standard Acceptance Criteria Re-activate Out of Service Equipment Q1.P1.P2(Con).P4(Con).P2(Con).P2(Con).P2(Con).P4(Con).P3(Con).P2(Con).P2(Con) Equal to Instructions Acknowledged
Equipment Problem Log (copy) | 11-Re-activate Out of Service Equipment P1.P2(Con).P4(Con).P2(Con).P2(Con).P3(Con).P2(Con) | Re-activate Out of Service Equipment Q1.P1.P2(Con).P4(Con).P2(Con).P2(Con).P3(Con).P2(Con) | Standard Acceptance Criteria Re-activate Out of Service Equipment Q1.P1.P2(Con).P4(Con).P2(Con).P2(Con).P3(Con).P2(Con) Equal to Instructions Acknowledged
Equipment Problem Log (copy) | 11-Re-activate Out of Service Equipment P1.P2(Con).P4(Con).P2(Con).P2(Con).P3(Con).P3(Con).P2(Con).P2(Con) | Re-activate Out of Service Equipment Q1.P1.P2(Con).P4(Con).P2(Con).P2(Con).P3(Con).P3(Con).P2(Con).P2(Con) | Standard Acceptance Criteria Re-activate Out of Service Equipment Q1.P1.P2(Con).P4(Con).P2(Con).P2(Con).P3(Con).P3(Con).P2(Con).P2(Con) Equal to Instructions Acknowledged
Equipment Problem Log (copy) | 11-Re-activate Out of Service Equipment P1.P2(Con).P4(Con).P3(Con).P2(Con).P2(Con).P2(Con) | Re-activate Out of Service Equipment Q1.P1.P2(Con).P4(Con).P3(Con).P2(Con).P2(Con).P2(Con) | Standard Acceptance Criteria Re-activate Out of Service Equipment Q1.P1.P2(Con).P4(Con).P3(Con).P2(Con).P2(Con).P2(Con) Equal to Instructions Acknowledged
Equipment Problem Log (copy) | 11-Re-activate Out of Service Equipment P1.P2(Con).P4(Con).P3(Con).P2(Con).P2(Con).P3(Con).P2(Con).P2(Con) | Re-activate Out of Service Equipment Q1.P1.P2(Con).P4(Con).P3(Con).P2(Con).P2(Con).P3(Con).P2(Con).P2(Con) | Standard Acceptance Criteria Re-activate Out of Service Equipment Q1.P1.P2(Con).P4(Con).P3(Con).P2(Con).P2(Con).P3(Con).P2(Con).P2(Con) Equal to Instructions Acknowledged
Equipment Problem Log (copy) | 11-Re-activate Out of Service Equipment P1.P2(Con).P4(Con).P3(Con).P2(Con).P3(Con).P2(Con).P2(Con) | Re-activate Out of Service Equipment Q1.P1.P2(Con).P4(Con).P3(Con).P2(Con).P3(Con).P2(Con).P2(Con) | Standard Acceptance Criteria Re-activate Out of Service Equipment Q1.P1.P2(Con).P4(Con).P3(Con).P2(Con).P3(Con).P2(Con).P2(Con) Equal to Instructions Acknowledged
Equipment Problem Log (copy) | 11-Re-activate Out of Service Equipment P1.P2(Con).P4(Con).P3(Con).P3(Con).P2(Con) | Re-activate Out of Service Equipment Q1.P1.P2(Con).P4(Con).P3(Con).P3(Con).P2(Con) | Standard Acceptance Criteria Re-activate Out of Service Equipment Q1.P1.P2(Con).P4(Con).P3(Con).P3(Con).P2(Con) Equal to Instructions Acknowledged
Equipment Problem Log (copy) | 11-Re-activate Out of Service Equipment P1.P2(Con).P4(Con).P3(Con).P3(Con).P3(Con).P2(Con).P2(Con) | Re-activate Out of Service Equipment Q1.P1.P2(Con).P4(Con).P3(Con).P3(Con).P3(Con).P2(Con).P2(Con) | Standard Acceptance Criteria Re-activate Out of Service Equipment Q1.P1.P2(Con).P4(Con).P3(Con).P3(Con).P3(Con).P2(Con).P2(Con) Equal to Instructions Acknowledged
Equipment Problem Log (copy) | 4-Service Call P1.P3(Con).P2(Con) | Problem Resolved? Q3.P1.P3(Con).P2(Con) | Standard Acceptance Criteria Problem Resolved? Q3.P1.P3(Con).P2(Con) Equal to Yes
Equipment Problem Log (copy) | 11-Re-activate Out of Service Equipment P1.P3(Con).P2(Con).P2(Con).P2(Con).P2(Con).P2(Con) | Re-activate Out of Service Equipment Q1.P1.P3(Con).P2(Con).P2(Con).P2(Con).P2(Con).P2(Con) | Standard Acceptance Criteria Re-activate Out of Service Equipment Q1.P1.P3(Con).P2(Con).P2(Con).P2(Con).P2(Con).P2(Con) Equal to Instructions Acknowledged
Equipment Problem Log (copy) | 11-Re-activate Out of Service Equipment P1.P3(Con).P2(Con).P2(Con).P2(Con).P2(Con).P3(Con).P2(Con).P2(Con) | Re-activate Out of Service Equipment Q1.P1.P3(Con).P2(Con).P2(Con).P2(Con).P2(Con).P3(Con).P2(Con).P2(Con) | Standard Acceptance Criteria Re-activate Out of Service Equipment Q1.P1.P3(Con).P2(Con).P2(Con).P2(Con).P2(Con).P3(Con).P2(Con).P2(Con) Equal to Instructions Acknowledged
Equipment Problem Log (copy) | 11-Re-activate Out of Service Equipment P1.P3(Con).P2(Con).P2(Con).P2(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P2(Con) | Re-activate Out of Service Equipment Q1.P1.P3(Con).P2(Con).P2(Con).P2(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P2(Con) | Standard Acceptance Criteria Re-activate Out of Service Equipment Q1.P1.P3(Con).P2(Con).P2(Con).P2(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P2(Con) Equal to Instructions Acknowledged
Equipment Problem Log (copy) | 11-Re-activate Out of Service Equipment P1.P3(Con).P2(Con).P2(Con).P2(Con).P3(Con).P2(Con).P2(Con) | Re-activate Out of Service Equipment Q1.P1.P3(Con).P2(Con).P2(Con).P2(Con).P3(Con).P2(Con).P2(Con) | Standard Acceptance Criteria Re-activate Out of Service Equipment Q1.P1.P3(Con).P2(Con).P2(Con).P2(Con).P3(Con).P2(Con).P2(Con) Equal to Instructions Acknowledged
Equipment Problem Log (copy) | 11-Re-activate Out of Service Equipment P1.P3(Con).P2(Con).P2(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P2(Con) | Re-activate Out of Service Equipment Q1.P1.P3(Con).P2(Con).P2(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P2(Con) | Standard Acceptance Criteria Re-activate Out of Service Equipment Q1.P1.P3(Con).P2(Con).P2(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P2(Con) Equal to Instructions Acknowledged
Equipment Problem Log (copy) | 11-Re-activate Out of Service Equipment P1.P3(Con).P2(Con).P2(Con).P3(Con).P2(Con) | Re-activate Out of Service Equipment Q1.P1.P3(Con).P2(Con).P2(Con).P3(Con).P2(Con) | Standard Acceptance Criteria Re-activate Out of Service Equipment Q1.P1.P3(Con).P2(Con).P2(Con).P3(Con).P2(Con) Equal to Instructions Acknowledged
Equipment Problem Log (copy) | 11-Re-activate Out of Service Equipment P1.P3(Con).P2(Con).P2(Con).P3(Con).P3(Con).P2(Con).P2(Con) | Re-activate Out of Service Equipment Q1.P1.P3(Con).P2(Con).P2(Con).P3(Con).P3(Con).P2(Con).P2(Con) | Standard Acceptance Criteria Re-activate Out of Service Equipment Q1.P1.P3(Con).P2(Con).P2(Con).P3(Con).P3(Con).P2(Con).P2(Con) Equal to Instructions Acknowledged
Equipment Problem Log (copy) | 11-Re-activate Out of Service Equipment P1.P3(Con).P2(Con).P2(Con).P3(Con).P3(Con).P2(Con).P3(Con).P2(Con).P2(Con) | Re-activate Out of Service Equipment Q1.P1.P3(Con).P2(Con).P2(Con).P3(Con).P3(Con).P2(Con).P3(Con).P2(Con).P2(Con) | Standard Acceptance Criteria Re-activate Out of Service Equipment Q1.P1.P3(Con).P2(Con).P2(Con).P3(Con).P3(Con).P2(Con).P3(Con).P2(Con).P2(Con) Equal to Instructions Acknowledged
Equipment Problem Log (copy) | 11-Re-activate Out of Service Equipment P1.P3(Con).P2(Con).P2(Con).P3(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P2(Con) | Re-activate Out of Service Equipment Q1.P1.P3(Con).P2(Con).P2(Con).P3(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P2(Con) | Standard Acceptance Criteria Re-activate Out of Service Equipment Q1.P1.P3(Con).P2(Con).P2(Con).P3(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P2(Con) Equal to Instructions Acknowledged
Equipment Problem Log (copy) | 11-Re-activate Out of Service Equipment P1.P3(Con).P2(Con).P2(Con).P3(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P2(Con) | Re-activate Out of Service Equipment Q1.P1.P3(Con).P2(Con).P2(Con).P3(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P2(Con) | Standard Acceptance Criteria Re-activate Out of Service Equipment Q1.P1.P3(Con).P2(Con).P2(Con).P3(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P2(Con) Equal to Instructions Acknowledged
Equipment Problem Log (copy) | 11-Re-activate Out of Service Equipment P1.P3(Con).P2(Con).P2(Con).P3(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P2(Con) | Re-activate Out of Service Equipment Q1.P1.P3(Con).P2(Con).P2(Con).P3(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P2(Con) | Standard Acceptance Criteria Re-activate Out of Service Equipment Q1.P1.P3(Con).P2(Con).P2(Con).P3(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P2(Con) Equal to Instructions Acknowledged
Equipment Problem Log (copy) | 11-Re-activate Out of Service Equipment P1.P3(Con).P2(Con).P2(Con).P3(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P2(Con) | Re-activate Out of Service Equipment Q1.P1.P3(Con).P2(Con).P2(Con).P3(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P2(Con) | Standard Acceptance Criteria Re-activate Out of Service Equipment Q1.P1.P3(Con).P2(Con).P2(Con).P3(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P2(Con) Equal to Instructions Acknowledged
Equipment Problem Log (copy) | 11-Re-activate Out of Service Equipment P1.P3(Con).P2(Con).P2(Con).P3(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P2(Con) | Re-activate Out of Service Equipment Q1.P1.P3(Con).P2(Con).P2(Con).P3(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P2(Con) | Standard Acceptance Criteria Re-activate Out of Service Equipment Q1.P1.P3(Con).P2(Con).P2(Con).P3(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P2(Con) Equal to Instructions Acknowledged
Equipment Problem Log (copy) | 11-Re-activate Out of Service Equipment P1.P3(Con).P2(Con).P2(Con).P3(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con) | Re-activate Out of Service Equipment Q1.P1.P3(Con).P2(Con).P2(Con).P3(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con) | Standard Acceptance Criteria Re-activate Out of Service Equipment Q1.P1.P3(Con).P2(Con).P2(Con).P3(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con) Equal to Instructions Acknowledged
Equipment Problem Log (copy) | 11-Re-activate Out of Service Equipment P1.P3(Con).P3(Con).P2(Con).P2(Con).P2(Con) | Re-activate Out of Service Equipment Q1.P1.P3(Con).P3(Con).P2(Con).P2(Con).P2(Con) | Standard Acceptance Criteria Re-activate Out of Service Equipment Q1.P1.P3(Con).P3(Con).P2(Con).P2(Con).P2(Con) Equal to Instructions Acknowledged
Equipment Problem Log (copy) | 11-Re-activate Out of Service Equipment P1.P3(Con).P3(Con).P2(Con).P2(Con).P3(Con).P2(Con).P2(Con) | Re-activate Out of Service Equipment Q1.P1.P3(Con).P3(Con).P2(Con).P2(Con).P3(Con).P2(Con).P2(Con) | Standard Acceptance Criteria Re-activate Out of Service Equipment Q1.P1.P3(Con).P3(Con).P2(Con).P2(Con).P3(Con).P2(Con).P2(Con) Equal to Instructions Acknowledged
Equipment Problem Log (copy) | 11-Re-activate Out of Service Equipment P1.P3(Con).P3(Con).P2(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P2(Con) | Re-activate Out of Service Equipment Q1.P1.P3(Con).P3(Con).P2(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P2(Con) | Standard Acceptance Criteria Re-activate Out of Service Equipment Q1.P1.P3(Con).P3(Con).P2(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P2(Con) Equal to Instructions Acknowledged
Equipment Problem Log (copy) | 11-Re-activate Out of Service Equipment P1.P3(Con).P3(Con).P2(Con).P3(Con).P2(Con).P2(Con) | Re-activate Out of Service Equipment Q1.P1.P3(Con).P3(Con).P2(Con).P3(Con).P2(Con).P2(Con) | Standard Acceptance Criteria Re-activate Out of Service Equipment Q1.P1.P3(Con).P3(Con).P2(Con).P3(Con).P2(Con).P2(Con) Equal to Instructions Acknowledged
Equipment Problem Log (copy) | 11-Re-activate Out of Service Equipment P1.P3(Con).P3(Con).P3(Con).P2(Con) | Re-activate Out of Service Equipment Q1.P1.P3(Con).P3(Con).P3(Con).P2(Con) | Standard Acceptance Criteria Re-activate Out of Service Equipment Q1.P1.P3(Con).P3(Con).P3(Con).P2(Con) Equal to Instructions Acknowledged
Equipment Problem Log (copy) | 11-Re-activate Out of Service Equipment P1.P3(Con).P3(Con).P3(Con).P3(Con).P2(Con).P2(Con) | Re-activate Out of Service Equipment Q1.P1.P3(Con).P3(Con).P3(Con).P3(Con).P2(Con).P2(Con) | Standard Acceptance Criteria Re-activate Out of Service Equipment Q1.P1.P3(Con).P3(Con).P3(Con).P3(Con).P2(Con).P2(Con) Equal to Instructions Acknowledged
Amicus Daily | Amicus Daily P1 | Start-Up Check Q1.P1 | Standard Acceptance Criteria Start-Up Check Q1.P1 Equal to Performed - OK
Amicus Daily | Amicus Daily P1 | Problem with Equipment Q2.P1 | Standard Acceptance Criteria Problem with Equipment Q2.P1 Equal to No
Spectra Optia Daily | Spectra Optia Daily P1 | Start-Up Check Q1.P1 | Standard Acceptance Criteria Start-Up Check Q1.P1 Equal to Performed - OK
Spectra Optia Daily | Spectra Optia Daily P1 | Problem with Equipment Q2.P1 | Standard Acceptance Criteria Problem with Equipment Q2.P1 Equal to No
Prism Shutdown Daily | Prism Shutdown Daily QC P1 | Run Purge on All Assigned Channels Q1.P1 | Standard Acceptance Criteria Run Purge on All Assigned Channels Q1.P1 Equal to Performed - OK
Prism Shutdown Daily | Prism Shutdown Daily QC P1 | Clean Activator Solution Bottle And Cap Q2.P1 | Standard Acceptance Criteria Clean Activator Solution Bottle And Cap Q2.P1 Equal to Performed - OK
Prism Shutdown Daily | Prism Shutdown Daily QC P1 | Clean Pipette Chute, Funnel Q3.P1 | Standard Acceptance Criteria Clean Pipette Chute, Funnel Q3.P1 Equal to Performed - OK
Prism Shutdown Daily | Prism Shutdown Daily QC P1 | Empty Pipette Tip and Reaction Tray Waste Containers Q4.P1 | Standard Acceptance Criteria Empty Pipette Tip and Reaction Tray Waste Containers Q4.P1 Equal to Performed - OK
Prism Shutdown Daily | Prism Shutdown Daily QC P1 | Problem with Equipment Q5.P1 | Standard Acceptance Criteria Problem with Equipment Q5.P1 Equal to No
Platelet Sampling Station Decontamination Daily | Platelet Sampling Station Decontamination Daily P1 | Platelet Sampling Station Decontamination Daily Q1.P1 | Standard Acceptance Criteria Platelet Sampling Station Decontamination Daily Q1.P1 Equal to Performed
Platelet Sampling Station Decontamination Daily | Platelet Sampling Station Decontamination Daily P1 | Problems with Equipment/QC Q2.P1 | Standard Acceptance Criteria Problems with Equipment/QC Q2.P1 Equal to No
Laboratory Developed External Qualification | Laboratory Developed External Qualification P1 | Test 1 Acceptable? Q7.P1 | Standard Acceptance Criteria Test 1 Acceptable? Q7.P1 Equal to Acceptable
Laboratory Developed External Qualification | Laboratory Developed External Qualification P1 | Test 2 Acceptable? Q12.P1 | Standard Acceptance Criteria Test 2 Acceptable? Q12.P1 Equal to Acceptable
Laboratory Developed External Qualification | Laboratory Developed External Qualification P1 | Test 3 Acceptable? Q17.P1 | Standard Acceptance Criteria Test 3 Acceptable? Q17.P1 Equal to Acceptable
Laboratory Developed External Qualification | Laboratory Developed External Qualification P1 | Test 4 Acceptable? Q22.P1 | Standard Acceptance Criteria Test 4 Acceptable? Q22.P1 Equal to Acceptable
Laboratory Developed External Qualification | Laboratory Developed External Qualification P1 | Test 5 Acceptable? Q27.P1 | Standard Acceptance Criteria Test 5 Acceptable? Q27.P1 Equal to Acceptable
Laboratory Developed External Qualification | Laboratory Developed External Qualification P1 | Test 6 Acceptable? Q32.P1 | Standard Acceptance Criteria Test 6 Acceptable? Q32.P1 Equal to Acceptable
Laboratory Developed External Qualification | Laboratory Developed External Qualification P1 | Test 7 Acceptable? Q37.P1 | Standard Acceptance Criteria Test 7 Acceptable? Q37.P1 Equal to Acceptable
Laboratory Developed External Qualification | Laboratory Developed External Qualification P1 | Test 8 Acceptable? Q42.P1 | Standard Acceptance Criteria Test 8 Acceptable? Q42.P1 Equal to Acceptable
Laboratory Developed External Qualification | Laboratory Developed External Qualification P1 | Test 9 Acceptable? Q47.P1 | Standard Acceptance Criteria Test 9 Acceptable? Q47.P1 Equal to Acceptable
Laboratory Developed External Qualification | Laboratory Developed External Qualification P1 | Test 10 Acceptable? Q52.P1 | Standard Acceptance Criteria Test 10 Acceptable? Q52.P1 Equal to Acceptable
Laboratory Developed External Qualification | Laboratory Developed External Qualification P1 | Test 11 Acceptable? Q57.P1 | Standard Acceptance Criteria Test 11 Acceptable? Q57.P1 Equal to Acceptable
Laboratory Developed External Qualification | Laboratory Developed External Qualification P1 | Test 12 Acceptable? Q62.P1 | Standard Acceptance Criteria Test 12 Acceptable? Q62.P1 Equal to Acceptable
Laboratory Developed External Qualification | Laboratory Developed External Qualification P1 | Test 13 Acceptable? Q67.P1 | Standard Acceptance Criteria Test 13 Acceptable? Q67.P1 Equal to Acceptable
Laboratory Developed External Qualification | Laboratory Developed External Qualification P1 | Test 14 Acceptable? Q72.P1 | Standard Acceptance Criteria Test 14 Acceptable? Q72.P1 Equal to Acceptable
Laboratory Developed External Qualification | Laboratory Developed External Qualification P1 | Test 15 Acceptable? Q77.P1 | Standard Acceptance Criteria Test 15 Acceptable? Q77.P1 Equal to Acceptable
Laboratory Developed External Qualification | Laboratory Developed External Qualification P1 | Problem with QC Q78.P1 | Standard Acceptance Criteria Problem with QC Q78.P1 Equal to No
Sickledex Lot Qualification As Needed | Sickledex Lot Qualification P3 | Current Sickledex Open Bottle Expiration Date Q1.P3 | Standard Acceptance Criteria Current Sickledex Open Bottle Expiration Date Q1.P3 Greater than or equal to Today
Sickledex Lot Qualification As Needed | Sickledex Lot Qualification P3 | New Sickledex Open Bottle Expiration Date Q2.P3 | Standard Acceptance Criteria New Sickledex Open Bottle Expiration Date Q2.P3 Greater than or equal to Today
Sickledex Lot Qualification As Needed | Sickledex Lot Qualification P3 | Current Sickledex + Positive Control Q3.P3 | Standard Acceptance Criteria Current Sickledex + Positive Control Q3.P3 Equal to Positive - Acceptable
Sickledex Lot Qualification As Needed | Sickledex Lot Qualification P3 | Current Sickledex + Negative Control Q4.P3 | Standard Acceptance Criteria Current Sickledex + Negative Control Q4.P3 Equal to Negative - Acceptable
Sickledex Lot Qualification As Needed | Sickledex Lot Qualification P3 | New Sickledex + Positive Control Q5.P3 | Standard Acceptance Criteria New Sickledex + Positive Control Q5.P3 Equal to Positive - Acceptable
Sickledex Lot Qualification As Needed | Sickledex Lot Qualification P3 | New Sickledex + Negative Control Q6.P3 | Standard Acceptance Criteria New Sickledex + Negative Control Q6.P3 Equal to Negative - Acceptable
Sickledex Lot Qualification As Needed | Sickledex Lot Qualification P3 | Sickledex New Lot Qualification Acceptable Q7.P3 | Standard Acceptance Criteria Sickledex New Lot Qualification Acceptable Q7.P3 Equal to Acceptable
Sickledex Lot Qualification As Needed | Sickledex Lot Qualification P3 | "This Lot Is Ready For Use" Green Sticker Placed on New Lot Q8.P3 | Standard Acceptance Criteria "This Lot Is Ready For Use" Green Sticker Placed on New Lot Q8.P3 Equal to Performed
Sickledex Lot Qualification As Needed | Sickledex Lot Qualification P3 | Problem with QC Q9.P3 | Standard Acceptance Criteria Problem with QC Q9.P3 Equal to No
Infectious Disease Reagent Qualification | Infectious Disease Reagent Qualification P2 | WNV Negative External Control Result Q2.P2 | Standard Acceptance Criteria WNV Negative External Control Result Q2.P2 Equal to Negative
Infectious Disease Reagent Qualification | Infectious Disease Reagent Qualification P2 | WNV Positive External Control Result Q3.P2 | Standard Acceptance Criteria WNV Positive External Control Result Q3.P2 Equal to Positive
Infectious Disease Reagent Qualification | Infectious Disease Reagent Qualification P2 | Ultrio Negative External Control Result Q4.P2 | Standard Acceptance Criteria Ultrio Negative External Control Result Q4.P2 Equal to Negative
Infectious Disease Reagent Qualification | Infectious Disease Reagent Qualification P2 | HIV-1 Positive External Control Result Q5.P2 | Standard Acceptance Criteria HIV-1 Positive External Control Result Q5.P2 Equal to Positive
Infectious Disease Reagent Qualification | Infectious Disease Reagent Qualification P2 | HCV Positive External Control Result Q6.P2 | Standard Acceptance Criteria HCV Positive External Control Result Q6.P2 Equal to Positive
Infectious Disease Reagent Qualification | Infectious Disease Reagent Qualification P2 | HBV Positive External Control Result Q7.P2 | Standard Acceptance Criteria HBV Positive External Control Result Q7.P2 Equal to Positive
Infectious Disease Reagent Qualification | Infectious Disease Reagent Qualification P2 | NAT Calibrators and External Controls Acceptable? Q8.P2 | Standard Acceptance Criteria NAT Calibrators and External Controls Acceptable? Q8.P2 Equal to Acceptable
Infectious Disease Reagent Qualification | Infectious Disease Reagent Qualification P2 | Prism Negative External Control Channel A Result Q10.P2 | Standard Acceptance Criteria Prism Negative External Control Channel A Result Q10.P2 Equal to Negative
Infectious Disease Reagent Qualification | Infectious Disease Reagent Qualification P2 | Prism Negative External Control Channel B Result Q11.P2 | Standard Acceptance Criteria Prism Negative External Control Channel B Result Q11.P2 Equal to Negative
Infectious Disease Reagent Qualification | Infectious Disease Reagent Qualification P2 | Prism Positive External Control Channel A Result Q12.P2 | Standard Acceptance Criteria Prism Positive External Control Channel A Result Q12.P2 Equal to Positive
Infectious Disease Reagent Qualification | Infectious Disease Reagent Qualification P2 | Prism Positive External Control Channel B Result Q13.P2 | Standard Acceptance Criteria Prism Positive External Control Channel B Result Q13.P2 Equal to Positive
Infectious Disease Reagent Qualification | Infectious Disease Reagent Qualification P2 | Prism Abbott Release Positive Channel A Result Q14.P2 | Standard Acceptance Criteria Prism Abbott Release Positive Channel A Result Q14.P2 Equal to Positive
Infectious Disease Reagent Qualification | Infectious Disease Reagent Qualification P2 | Prism Abbott Release Positive Channel B Result Q15.P2 | Standard Acceptance Criteria Prism Abbott Release Positive Channel B Result Q15.P2 Equal to Positive
Infectious Disease Reagent Qualification | Infectious Disease Reagent Qualification P2 | Prism Calibrators and External Controls Acceptable? Q16.P2 | Standard Acceptance Criteria Prism Calibrators and External Controls Acceptable? Q16.P2 Equal to Acceptable
Infectious Disease Reagent Qualification | Infectious Disease Reagent Qualification P2 | Problem with Equipment Q17.P2 | Standard Acceptance Criteria Problem with Equipment Q17.P2 Equal to No
Assay Incident Report As Needed | Assay Incident Report P1 | Date of Occurrence: Q3.P1 | Standard Acceptance Criteria Date of Occurrence: Q3.P1 Less than or equal to Today
Assay Incident Report As Needed | Assay Incident Report P1 | Problem with Equipment Q6.P1 | Standard Acceptance Criteria Problem with Equipment Q6.P1 Equal to No
Flow Cytometry Reagent Qualification | Flow Cytometry Reagent Qualificiation P2 | Controls Acceptable? Q3.P2 | Standard Acceptance Criteria Controls Acceptable? Q3.P2 Equal to Acceptable
Flow Cytometry Reagent Qualification | Flow Cytometry Reagent Qualificiation P2 | Controls Acceptable? Q6.P2 | Standard Acceptance Criteria Controls Acceptable? Q6.P2 Equal to Acceptable
Flow Cytometry Reagent Qualification | Flow Cytometry Reagent Qualificiation P2 | Problem with Equipment Q7.P2 | Standard Acceptance Criteria Problem with Equipment Q7.P2 Equal to No
Gammacell 3000 Irradiator Bi-Annual | Gammacell 3000 Irradiator Biannual P1 | Leak Test/Check Performed? Q1.P1 | Standard Acceptance Criteria Leak Test/Check Performed? Q1.P1 Equal to Performed - OK
Gammacell 3000 Irradiator Bi-Annual | Gammacell 3000 Irradiator Biannual P1 | Electrical Check Performed? Q2.P1 | Standard Acceptance Criteria Electrical Check Performed? Q2.P1 Equal to Performed - OK
IBL 437 Irradiator Bi-Annual | IBL 437 Irradiator Bi Annual Maintenance P1 | Leak Test/Check Performed? Q1.P1 | Standard Acceptance Criteria Leak Test/Check Performed? Q1.P1 Equal to Performed - OK
IBL 437 Irradiator Bi-Annual | IBL 437 Irradiator Bi Annual Maintenance P1 | Electrical Check Performed? Q2.P1 | Standard Acceptance Criteria Electrical Check Performed? Q2.P1 Equal to Performed - OK
Digital Timer Calibration Annual | Annual Digital Timer Calibration P1 | Date Sent Out Q1.P1 | Standard Acceptance Criteria Date Sent Out Q1.P1 Less than or equal to Today
CM Electronic Balance Monthly | CM Electronic Balance Monthly P1 | Readout (g) After Tared Q1.P1 | Standard Acceptance Criteria Readout (g) After Tared Q1.P1 Between -0.01 to 0.01
CM Electronic Balance Monthly | CM Electronic Balance Monthly P1 | Readout 5 G Weight Q2.P1 | Standard Acceptance Criteria Readout 5 G Weight Q2.P1 Between 4.9 to 5.1
CM Electronic Balance Monthly | CM Electronic Balance Monthly P1 | Readout 30 G Weight Q3.P1 | Standard Acceptance Criteria Readout 30 G Weight Q3.P1 Between 29.9 to 30.1
CM Electronic Balance Monthly | CM Electronic Balance Monthly P1 | Readout 100 G Weight Q4.P1 | Standard Acceptance Criteria Readout 100 G Weight Q4.P1 Between 99.9 to 100.1
CM Electronic Balance Monthly | CM Electronic Balance Monthly P1 | Readout 300 G Weight Q5.P1 | Standard Acceptance Criteria Readout 300 G Weight Q5.P1 Between 299.9 to 300.1
CM Electronic Balance Monthly | CM Electronic Balance Monthly P1 | Readout 500 G Weight Q6.P1 | Standard Acceptance Criteria Readout 500 G Weight Q6.P1 Between 499.9 to 500.1
CM Electronic Balance Monthly | CM Electronic Balance Monthly P1 | Readout 1000 G Weight Q7.P1 | Standard Acceptance Criteria Readout 1000 G Weight Q7.P1 Between 999.9 to 1000.1
CM Electronic Balance Monthly | CM Electronic Balance Monthly P1 | Problem with Equipment Q8.P1 | Standard Acceptance Criteria Problem with Equipment Q8.P1 Equal to No
Fixed Sites Electronic Balance Monthly | Electronic Balance DO Monthly P1 | Readout (g) After Tared Q1.P1 | Standard Acceptance Criteria Readout (g) After Tared Q1.P1 Between -0.01 to 0.01
Fixed Sites Electronic Balance Monthly | Electronic Balance DO Monthly P1 | Readout 300 G Weight Q2.P1 | Standard Acceptance Criteria Readout 300 G Weight Q2.P1 Between 299.9 to 300.1
Fixed Sites Electronic Balance Monthly | Electronic Balance DO Monthly P1 | Readout 500 G Weight Q3.P1 | Standard Acceptance Criteria Readout 500 G Weight Q3.P1 Between 499.9 to 500.1
Fixed Sites Electronic Balance Monthly | Electronic Balance DO Monthly P1 | Readout 1000 G Weight Q4.P1 | Standard Acceptance Criteria Readout 1000 G Weight Q4.P1 Between 999.9 to 1000.1
Fixed Sites Electronic Balance Monthly | Electronic Balance DO Monthly P1 | Problem with Equipment Q5.P1 | Standard Acceptance Criteria Problem with Equipment Q5.P1 Equal to No
Irradiator Dosimetry Annual | Problem with Equipment P2 | Problem with Equipment Q1.P2 | Standard Acceptance Criteria Problem with Equipment Q1.P2 Equal to No
Coulter AcT Diff 2 Daily Duplicate | Coulter Lot P1 | Normal Control Open Bottle Expiration Date Q2.P1 | Standard Acceptance Criteria Normal Control Open Bottle Expiration Date Q2.P1 Greater than or equal to Today
Coulter AcT Diff 2 Daily Duplicate | Coulter Lot P1 | High Control Open Bottle Expiration Date Q4.P1 | Standard Acceptance Criteria High Control Open Bottle Expiration Date Q4.P1 Greater than or equal to Today
Coulter AcT Diff 2 Daily Duplicate | Startup Check P2 | Start-Up Check Q1.P2 | Standard Acceptance Criteria Start-Up Check Q1.P2 Equal to Performed - OK
Coulter AcT Diff 2 Daily Duplicate | Coulter Daily P3 | Normal Control Value-WBC Q11.P3 | Standard Acceptance Criteria Normal Control Value-WBC Q11.P3 Between 9.3 to 10.2
Coulter AcT Diff 2 Daily Duplicate | Coulter Daily P3 | Normal Control Value - RBC Q1.P3 | Standard Acceptance Criteria Normal Control Value - RBC Q1.P3 Between 3.84 to 4.2
Coulter AcT Diff 2 Daily Duplicate | Coulter Daily P3 | Normal Control Value - HGB Q2.P3 | Standard Acceptance Criteria Normal Control Value - HGB Q2.P3 Between 11 to 12.2
Coulter AcT Diff 2 Daily Duplicate | Coulter Daily P3 | Normal Control Value - HCT Q3.P3 | Standard Acceptance Criteria Normal Control Value - HCT Q3.P3 Between 31.2 to 34.8
Coulter AcT Diff 2 Daily Duplicate | Coulter Daily P3 | Normal Control Value PLT Q4.P3 | Standard Acceptance Criteria Normal Control Value PLT Q4.P3 Between 176 to 254
Coulter AcT Diff 2 Daily Duplicate | Coulter Daily P3 | Abnormal High Control Value - WBC Q5.P3 | Standard Acceptance Criteria Abnormal High Control Value - WBC Q5.P3 Between 17.8 to 19.6
Coulter AcT Diff 2 Daily Duplicate | Coulter Daily P3 | Abnormal High Control Value - RBC Q6.P3 | Standard Acceptance Criteria Abnormal High Control Value - RBC Q6.P3 Between 4.97 to 5.45
Coulter AcT Diff 2 Daily Duplicate | Coulter Daily P3 | Abnormal High Control Value - HGB Q7.P3 | Standard Acceptance Criteria Abnormal High Control Value - HGB Q7.P3 Between 16.1 to 17.3
Coulter AcT Diff 2 Daily Duplicate | Coulter Daily P3 | Abnormal High Control Value - HCT Q8.P3 | Standard Acceptance Criteria Abnormal High Control Value - HCT Q8.P3 Between 44.9 to 49.7
Coulter AcT Diff 2 Daily Duplicate | Coulter Daily P3 | Abnormal High Control Value - PLT Q9.P3 | Standard Acceptance Criteria Abnormal High Control Value - PLT Q9.P3 Between 319 to 427
Coulter AcT Diff 2 Daily Duplicate | Coulter Daily P3 | Clean and Shut Down Instrument? Q10.P3 | Standard Acceptance Criteria Clean and Shut Down Instrument? Q10.P3 Equal to Performed - OK
Coulter AcT Diff 2 Daily Duplicate | Coulter Daily P3 | Problem with Equipment Q12.P3 | Standard Acceptance Criteria Problem with Equipment Q12.P3 Equal to No
Coulter AcT Diff 2 Daily Duplicate | Shutdown P4 | Clean and Shut Down Instrument? Q1.P4 | Standard Acceptance Criteria Clean and Shut Down Instrument? Q1.P4 Equal to Performed - OK
Coulter AcT Diff 2 Daily Duplicate | Shutdown P4 | Problem with Equipment Q2.P4 | Standard Acceptance Criteria Problem with Equipment Q2.P4 Equal to No
Trip Scale Daily | Trip Scale Daily P2 | Calibration Passed? Q2.P2 | Standard Acceptance Criteria Calibration Passed? Q2.P2 Equal to Passed
Trip Scale Daily | Trip Scale Daily P2 | End of Day Cleaning Q1.P2 | Standard Acceptance Criteria End of Day Cleaning Q1.P2 Equal to Performed
Trip Scale Daily | Trip Scale Daily P2 | Problem with Equipment Q3.P2 | Standard Acceptance Criteria Problem with Equipment Q3.P2 Equal to No
Electronic Balance PM Annual | PM P1 | PM Performed Q3.P1 | Standard Acceptance Criteria PM Performed Q3.P1 Equal to Performed - OK
Electronic Balance PM Annual | PM P1 | Problem with Equipment Q4.P1 | Standard Acceptance Criteria Problem with Equipment Q4.P1 Equal to No
Table Top Sealer Weekly | Table Top Sealer Weekly P1 | Clean/Wipe the Heat Sealing Element Q1.P1 | Standard Acceptance Criteria Clean/Wipe the Heat Sealing Element Q1.P1 Equal to Performed - OK
Table Top Sealer Weekly | Table Top Sealer Weekly P1 | Sealer Disinfected Q2.P1 | Standard Acceptance Criteria Sealer Disinfected Q2.P1 Equal to Performed
Table Top Sealer Weekly | Table Top Sealer Weekly P1 | Problem with Equipment Q3.P1 | Standard Acceptance Criteria Problem with Equipment Q3.P1 Equal to No
Spectra Cobe Weekly | Spectra Cobe Weekly P1 | Inlet Air Detector Cleaned Q1.P1 | Standard Acceptance Criteria Inlet Air Detector Cleaned Q1.P1 Equal to Performed
Spectra Cobe Weekly | Spectra Cobe Weekly P1 | Return Air Detector Cleaned Q2.P1 | Standard Acceptance Criteria Return Air Detector Cleaned Q2.P1 Equal to Performed
Spectra Cobe Weekly | Spectra Cobe Weekly P1 | RBC Detector Cleaned Q3.P1 | Standard Acceptance Criteria RBC Detector Cleaned Q3.P1 Equal to Performed
Spectra Cobe Weekly | Spectra Cobe Weekly P1 | Fluid Leak Detector Cleaned Q4.P1 | Standard Acceptance Criteria Fluid Leak Detector Cleaned Q4.P1 Equal to Performed
Spectra Cobe Weekly | Spectra Cobe Weekly P1 | Centrifuge Pressure Sensor Cleaned Q5.P1 | Standard Acceptance Criteria Centrifuge Pressure Sensor Cleaned Q5.P1 Equal to Performed
Spectra Cobe Weekly | Spectra Cobe Weekly P1 | Collection Concentration Monitor (CCM) Cleaned Q6.P1 | Standard Acceptance Criteria Collection Concentration Monitor (CCM) Cleaned Q6.P1 Equal to Performed
Spectra Cobe Weekly | Spectra Cobe Weekly P1 | Door Tracks Cleaned Q7.P1 | Standard Acceptance Criteria Door Tracks Cleaned Q7.P1 Equal to Performed
Spectra Cobe Weekly | Spectra Cobe Weekly P1 | Problem with Equipment Q8.P1 | Standard Acceptance Criteria Problem with Equipment Q8.P1 Equal to No
Vacutainer Tubes Daily | Problem with QC P1 | Problem with QC Q1.P1 | Standard Acceptance Criteria Problem with QC Q1.P1 Equal to No
HCG Pregnancy Daily | QC Performed P1 | Was the QC Performed Q1.P1 | Standard Acceptance Criteria Was the QC Performed Q1.P1 Equal to No
HCG Pregnancy Daily | Biorad Lot P1.P2(Con) | Biorad Liquichek Urine Chemistry Control Level 1 Open Bottle Expiration Date Q1.P1.P2(Con) | Standard Acceptance Criteria Biorad Liquichek Urine Chemistry Control Level 1 Open Bottle Expiration Date Q1.P1.P2(Con) Greater than or equal to Today
HCG Pregnancy Daily | Biorad Lot P1.P2(Con) | Biorad Liquichek Urine Chemistry Control Level 2 Open Bottle Expiration Date Q2.P1.P2(Con) | Standard Acceptance Criteria Biorad Liquichek Urine Chemistry Control Level 2 Open Bottle Expiration Date Q2.P1.P2(Con) Greater than or equal to Today
HCG Pregnancy Daily | HCG Pregancy Daily P1.P3(Con) | Level 1 Control Result Q2.P1.P3(Con) | Standard Acceptance Criteria Level 1 Control Result Q2.P1.P3(Con) Equal to Negative
HCG Pregnancy Daily | HCG Pregancy Daily P1.P3(Con) | Level 2 Control Result Q1.P1.P3(Con) | Standard Acceptance Criteria Level 2 Control Result Q1.P1.P3(Con) Equal to Positive
HCG Pregnancy Daily | Problem with QC P1.P4(Con) | Problem with QC Q1.P1.P4(Con) | Standard Acceptance Criteria Problem with QC Q1.P1.P4(Con) Equal to No
Irradiator PM Annual | PM P1 | PM Performed Q3.P1 | Standard Acceptance Criteria PM Performed Q3.P1 Equal to Performed - OK
Irradiator PM Annual | PM P1 | Problem with Equipment Q4.P1 | Standard Acceptance Criteria Problem with Equipment Q4.P1 Equal to No
Trima As Needed | Blood Spill P1 | Blood Spill Occurred Q1.P1 | Standard Acceptance Criteria Blood Spill Occurred Q1.P1 Equal to No
Trima As Needed | Trima As Needed P2 | Trima System's Surface Disinfected Q1.P2 | Standard Acceptance Criteria Trima System's Surface Disinfected Q1.P2 Equal to Yes
Trima As Needed | Trima As Needed P2 | Problem with Equipment Q2.P2 | Standard Acceptance Criteria Problem with Equipment Q2.P2 Equal to No
Hemoflow 400 Validation As Needed | Hemoflow 400 As Needed Validation P1 | Hemoflow System OK? Q2.P1 | Standard Acceptance Criteria Hemoflow System OK? Q2.P1 Equal to Yes
Hemoflow 400 Validation As Needed | Hemoflow 400 As Needed Validation P1 | Record 200 G Value Displayed on Hemoflow (g) Q3.P1 | Standard Acceptance Criteria Record 200 G Value Displayed on Hemoflow (g) Q3.P1 Between 198 to 202
Hemoflow 400 Validation As Needed | Hemoflow 400 As Needed Validation P1 | Record 500 G Value Displayed on Hemoflow (g) Q4.P1 | Standard Acceptance Criteria Record 500 G Value Displayed on Hemoflow (g) Q4.P1 Between 495 to 505
Hemoflow 400 Validation As Needed | Hemoflow 400 As Needed Validation P1 | Record 700 G Value Displayed on Hemoflow (g) Q5.P1 | Standard Acceptance Criteria Record 700 G Value Displayed on Hemoflow (g) Q5.P1 Between 693 to 707
Hemoflow 400 Validation As Needed | Hemoflow 400 As Needed Validation P1 | Record 1000 G Value Displayed on Hemoflow (g) Q6.P1 | Standard Acceptance Criteria Record 1000 G Value Displayed on Hemoflow (g) Q6.P1 Between 990 to 1010
Hemoflow 400 Validation As Needed | Hemoflow 400 As Needed Validation P1 | Clamp Operation Test Acceptable Q7.P1 | Standard Acceptance Criteria Clamp Operation Test Acceptable Q7.P1 Equal to Acceptable
Hemoflow 400 Validation As Needed | Hemoflow 400 As Needed Validation P1 | Record Leak Test at 30 Seconds Q8.P1 | Standard Acceptance Criteria Record Leak Test at 30 Seconds Q8.P1 Between 28 to 32
Hemoflow 400 Validation As Needed | Hemoflow 400 As Needed Validation P1 | Mixer Test Acceptable Q9.P1 | Standard Acceptance Criteria Mixer Test Acceptable Q9.P1 Equal to Acceptable
Hemoflow 400 Validation As Needed | Hemoflow 400 As Needed Validation P1 | Date Displayed on Hemoflow (should Be Current Date) Q10.P1 | Standard Acceptance Criteria Date Displayed on Hemoflow (should Be Current Date) Q10.P1 Less than or equal to Today
Hemoflow 400 Validation As Needed | Hemoflow 400 As Needed Validation P1 | Hemoflow Validation Results Acceptable Q19.P1 | Standard Acceptance Criteria Hemoflow Validation Results Acceptable Q19.P1 Equal to Acceptable
Hemoflow 400 Validation As Needed | Hemoflow 400 As Needed Validation P1 | Problem with Equipment Q20.P1 | Standard Acceptance Criteria Problem with Equipment Q20.P1 Equal to No
HCG Pregnancy Kit Lot-to-Lot | HCG Pregnancy Lot to Lot Info P1 | Level 1 Control Open Bottle Expiration Date Q1.P1 | Standard Acceptance Criteria Level 1 Control Open Bottle Expiration Date Q1.P1 Greater than or equal to Today
HCG Pregnancy Kit Lot-to-Lot | HCG Pregnancy Lot to Lot Info P1 | Level 2 Control Open Bottle Expiration Date Q2.P1 | Standard Acceptance Criteria Level 2 Control Open Bottle Expiration Date Q2.P1 Greater than or equal to Today
HCG Pregnancy Kit Lot-to-Lot | HCG Pregnancy Lot to Lot Info P1 | Current Kit Level 1 Internal Control Line Present Q3.P1 | Standard Acceptance Criteria Current Kit Level 1 Internal Control Line Present Q3.P1 Equal to Yes
HCG Pregnancy Kit Lot-to-Lot | HCG Pregnancy Lot to Lot Info P1 | New Kit Level 1 Internal Control Line Present Q4.P1 | Standard Acceptance Criteria New Kit Level 1 Internal Control Line Present Q4.P1 Equal to Yes
HCG Pregnancy Kit Lot-to-Lot | HCG Pregnancy Lot to Lot Info P1 | Current Kit Level 1 Result Q5.P1 | Standard Acceptance Criteria Current Kit Level 1 Result Q5.P1 Equal to Negative:Acceptable
HCG Pregnancy Kit Lot-to-Lot | HCG Pregnancy Lot to Lot Info P1 | New Kit Level 1 Result Q6.P1 | Standard Acceptance Criteria New Kit Level 1 Result Q6.P1 Equal to Negative:Acceptable
HCG Pregnancy Kit Lot-to-Lot | HCG Pregnancy Lot to Lot Info P1 | Current Kit Level 2 Internal Control Line Present Q7.P1 | Standard Acceptance Criteria Current Kit Level 2 Internal Control Line Present Q7.P1 Equal to Yes
HCG Pregnancy Kit Lot-to-Lot | HCG Pregnancy Lot to Lot Info P1 | New Kit Level 2 Internal Control Line Present Q8.P1 | Standard Acceptance Criteria New Kit Level 2 Internal Control Line Present Q8.P1 Equal to Yes
HCG Pregnancy Kit Lot-to-Lot | HCG Pregnancy Lot to Lot Info P1 | Current Kit Level 2 Result Q9.P1 | Standard Acceptance Criteria Current Kit Level 2 Result Q9.P1 Equal to Positive:Acceptable
HCG Pregnancy Kit Lot-to-Lot | HCG Pregnancy Lot to Lot Info P1 | New Kit Level 2 Result Q10.P1 | Standard Acceptance Criteria New Kit Level 2 Result Q10.P1 Equal to Positive:Acceptable
HCG Pregnancy Kit Lot-to-Lot | HCG Pregnancy Lot to Lot Info P1 | New Kit Level 1 Acceptable for Lot Qualification Q11.P1 | Standard Acceptance Criteria New Kit Level 1 Acceptable for Lot Qualification Q11.P1 Equal to Yes
HCG Pregnancy Kit Lot-to-Lot | HCG Pregnancy Lot to Lot Info P1 | New Kit Level 2 Acceptable for Lot Qualification Q12.P1 | Standard Acceptance Criteria New Kit Level 2 Acceptable for Lot Qualification Q12.P1 Equal to Yes
HCG Pregnancy Kit Lot-to-Lot | HCG Pregnancy Lot to Lot Info P1 | If Both Lots Are Acceptable, Mark New Lot with "This Lot Is Ready For Use" Green Sticker Q13.P1 | Standard Acceptance Criteria If Both Lots Are Acceptable, Mark New Lot with "This Lot Is Ready For Use" Green Sticker Q13.P1 Equal to Performed
HCG Pregnancy Kit Lot-to-Lot | HCG Pregnancy Lot to Lot Info P1 | Problem with QC Q14.P1 | Standard Acceptance Criteria Problem with QC Q14.P1 Equal to No
Amicus As Needed | Amicus As Needed P1 | Amicus Separator Cleaned Q1.P1 | Standard Acceptance Criteria Amicus Separator Cleaned Q1.P1 Equal to Yes
Amicus As Needed | Amicus As Needed P1 | Pump Cleaned Q2.P1 | Standard Acceptance Criteria Pump Cleaned Q2.P1 Equal to Yes
Amicus As Needed | Amicus As Needed P1 | Pump Head Recess Cleaned Q3.P1 | Standard Acceptance Criteria Pump Head Recess Cleaned Q3.P1 Equal to Yes
Amicus As Needed | Amicus As Needed P1 | Air Detector Cleaned Q4.P1 | Standard Acceptance Criteria Air Detector Cleaned Q4.P1 Equal to Yes
Amicus As Needed | Amicus As Needed P1 | Optical Sensor Cleaned Q5.P1 | Standard Acceptance Criteria Optical Sensor Cleaned Q5.P1 Equal to Yes
Amicus As Needed | Amicus As Needed P1 | Cassette Plate Cleaned Q6.P1 | Standard Acceptance Criteria Cassette Plate Cleaned Q6.P1 Equal to Yes
Amicus As Needed | Amicus As Needed P1 | Spool Cleaned Q7.P1 | Standard Acceptance Criteria Spool Cleaned Q7.P1 Equal to Yes
Amicus As Needed | Amicus As Needed P1 | Spool Holder Cleaned Q8.P1 | Standard Acceptance Criteria Spool Holder Cleaned Q8.P1 Equal to Yes
Amicus As Needed | Amicus As Needed P1 | Centrifuge Cleaned Q9.P1 | Standard Acceptance Criteria Centrifuge Cleaned Q9.P1 Equal to Yes
Amicus As Needed | Amicus As Needed P1 | Centrifuge Compartment Cleaned Q10.P1 | Standard Acceptance Criteria Centrifuge Compartment Cleaned Q10.P1 Equal to Yes
Amicus As Needed | Amicus As Needed P1 | Problem with Equipment Q12.P1 | Standard Acceptance Criteria Problem with Equipment Q12.P1 Equal to No
Spectra Optia As Needed | Optia As Needed P2 | Sensors Cleaned Q1.P2 | Standard Acceptance Criteria Sensors Cleaned Q1.P2 Equal to Performed - OK
Spectra Optia As Needed | Optia As Needed P2 | Detectors Cleaned Q2.P2 | Standard Acceptance Criteria Detectors Cleaned Q2.P2 Equal to Performed - OK
Spectra Optia As Needed | Optia As Needed P2 | Pump Housing Cleaned Q3.P2 | Standard Acceptance Criteria Pump Housing Cleaned Q3.P2 Equal to Performed
Spectra Optia As Needed | Optia As Needed P2 | Pump Rotors Cleaned Q4.P2 | Standard Acceptance Criteria Pump Rotors Cleaned Q4.P2 Equal to Performed
Spectra Optia As Needed | Optia As Needed P2 | Fluid Leak Detector Cleaned Q5.P2 | Standard Acceptance Criteria Fluid Leak Detector Cleaned Q5.P2 Equal to Performed - OK
Spectra Optia As Needed | Optia As Needed P2 | Glass Covers on Light Cleaned Q6.P2 | Standard Acceptance Criteria Glass Covers on Light Cleaned Q6.P2 Equal to Performed
Spectra Optia As Needed | Optia As Needed P2 | Equipment Disinfected After Blood Spill? Q7.P2 | Standard Acceptance Criteria Equipment Disinfected After Blood Spill? Q7.P2 Equal to Performed
Spectra Optia As Needed | Optia As Needed P2 | Problem with Equipment Q9.P2 | Standard Acceptance Criteria Problem with Equipment Q9.P2 Equal to No
Hemocue Lot to Lot | Hemocue Lot to Lot Information P1 | Current Master Lot Manufacturer Expiration Date Q2.P1 | Standard Acceptance Criteria Current Master Lot Manufacturer Expiration Date Q2.P1 Greater than or equal to Today
Hemocue Lot to Lot | Hemocue Lot to Lot Information P1 | Current Normal Control Manufacturer Expiration Date Q4.P1 | Standard Acceptance Criteria Current Normal Control Manufacturer Expiration Date Q4.P1 Greater than or equal to Today
Hemocue Lot to Lot | Hemocue Lot to Lot Information P1 | Current Abnnormal High Control Manufacturer Expiration Date Q6.P1 | Standard Acceptance Criteria Current Abnnormal High Control Manufacturer Expiration Date Q6.P1 Greater than or equal to Today
Hemocue Lot to Lot | Hemocue Lot to Lot Information P1 | New Master Lot Manufacturer Expiration Date Q8.P1 | Standard Acceptance Criteria New Master Lot Manufacturer Expiration Date Q8.P1 Greater than or equal to Today
Hemocue Lot to Lot | Hemocue Lot to Lot Information P1 | New Normal Control Manufacturer Expiration Date Q10.P1 | Standard Acceptance Criteria New Normal Control Manufacturer Expiration Date Q10.P1 Greater than or equal to Today
Hemocue Lot to Lot | Hemocue Lot to Lot Information P1 | New Abnnormal High Control Manufacturer Expiration Date Q12.P1 | Standard Acceptance Criteria New Abnnormal High Control Manufacturer Expiration Date Q12.P1 Greater than or equal to Today
Hemocue Lot to Lot | Hemocue Lot to Lot P2 | Current Normal Control Open Bottle Expiration Date Q2.P2 | Standard Acceptance Criteria Current Normal Control Open Bottle Expiration Date Q2.P2 Greater than or equal to Today
Hemocue Lot to Lot | Hemocue Lot to Lot P2 | Current Abnormal High Control Open Bottle Expiration Date Q3.P2 | Standard Acceptance Criteria Current Abnormal High Control Open Bottle Expiration Date Q3.P2 Greater than or equal to Today
Hemocue Lot to Lot | Hemocue Lot to Lot P2 | Accept Current Lot Levy-Jennings Chart? Q6.P2 | Standard Acceptance Criteria Accept Current Lot Levy-Jennings Chart? Q6.P2 Equal to Acceptable
Hemocue Lot to Lot | Hemocue Lot to Lot P2 | New Normal Control Open Bottle Expiration Date Q7.P2 | Standard Acceptance Criteria New Normal Control Open Bottle Expiration Date Q7.P2 Greater than or equal to Today
Hemocue Lot to Lot | Hemocue Lot to Lot P2 | New Abnormal High Control Open Bottle Expiration Date Q8.P2 | Standard Acceptance Criteria New Abnormal High Control Open Bottle Expiration Date Q8.P2 Greater than or equal to Today
Hemocue Lot to Lot | Hemocue Lot to Lot P2 | Accept Current Lot Levy-Jennings Chart? Q11.P2 | Standard Acceptance Criteria Accept Current Lot Levy-Jennings Chart? Q11.P2 Equal to Acceptable
Hemocue Lot to Lot | Equipment Problem Log P2.P2(Con) | Date of Problem Q1.P2.P2(Con) | Standard Acceptance Criteria Date of Problem Q1.P2.P2(Con) Less than or equal to Today
Hemocue Lot to Lot | Equipment Problem Log P2.P2(Con) | Date of Trouble Shooting Q8.P2.P2(Con) | Standard Acceptance Criteria Date of Trouble Shooting Q8.P2.P2(Con) Less than or equal to Today
Hemocue Lot to Lot | Equipment Problem Log P2.P2(Con) | Problem Resolved? Q10.P2.P2(Con) | Standard Acceptance Criteria Problem Resolved? Q10.P2.P2(Con) Equal to Yes
Hemocue Lot to Lot | Equipment Problem Log P2.P2(Con) | Remove from Use? Q11.P2.P2(Con) | Standard Acceptance Criteria Remove from Use? Q11.P2.P2(Con) Equal to No
Hemocue Lot to Lot | Equipment Problem Log P2.P2(Con) | Service Called? Q14.P2.P2(Con) | Standard Acceptance Criteria Service Called? Q14.P2.P2(Con) Equal to Yes
Hemocue Lot to Lot | Equipment Problem Log P2.P2(Con) | Date Service Called Q15.P2.P2(Con) | Standard Acceptance Criteria Date Service Called Q15.P2.P2(Con) Less than or equal to Today
Hemocue Lot to Lot | Equipment Problem Log P2.P2(Con) | Equipment Removed? Q18.P2.P2(Con) | Standard Acceptance Criteria Equipment Removed? Q18.P2.P2(Con) Equal to No
Hemocue Lot to Lot | Equipment Problem Log P2.P2(Con) | Date Removed Q19.P2.P2(Con) | Standard Acceptance Criteria Date Removed Q19.P2.P2(Con) Less than or equal to Today
Hemocue Lot to Lot | Equipment Problem Log P2.P2(Con) | Date Returned Q20.P2.P2(Con) | Standard Acceptance Criteria Date Returned Q20.P2.P2(Con) Less than or equal to Today
Hemocue Lot to Lot | Equipment Problem Log P2.P2(Con) | Qualification Validation Required? Q21.P2.P2(Con) | Standard Acceptance Criteria Qualification Validation Required? Q21.P2.P2(Con) Equal to Yes
Hemocue Lot to Lot | Equipment Problem Log P2.P2(Con) | Date Qualification Completed Q22.P2.P2(Con) | Standard Acceptance Criteria Date Qualification Completed Q22.P2.P2(Con) Less than or equal to Today
Hemocue Lot to Lot | Equipment Problem Log P2.P2(Con) | Qualification/Validation Notes Q23.P2.P2(Con) | Standard Acceptance Criteria Qualification/Validation Notes Q23.P2.P2(Con) undefined undefined
Hemocue Lot to Lot | Equipment Problem Log P2.P2(Con) | Qualification/Validation Acceptable? Q24.P2.P2(Con) | Standard Acceptance Criteria Qualification/Validation Acceptable? Q24.P2.P2(Con) Equal to Yes
Hemocue Lot to Lot | Equipment Problem Log P2.P2(Con) | Date Approved Fo Use Q25.P2.P2(Con) | Standard Acceptance Criteria Date Approved Fo Use Q25.P2.P2(Con) Less than or equal to Today
Hemocue Lot to Lot | Equipment Problem Log P2.P3(Con) | Date of Problem Q1.P2.P3(Con) | Standard Acceptance Criteria Date of Problem Q1.P2.P3(Con) Less than or equal to Today
Hemocue Lot to Lot | Equipment Problem Log P2.P3(Con) | Date of Trouble Shooting Q8.P2.P3(Con) | Standard Acceptance Criteria Date of Trouble Shooting Q8.P2.P3(Con) Less than or equal to Today
Hemocue Lot to Lot | Equipment Problem Log P2.P3(Con) | Problem Resolved? Q10.P2.P3(Con) | Standard Acceptance Criteria Problem Resolved? Q10.P2.P3(Con) Equal to Yes
Hemocue Lot to Lot | Equipment Problem Log P2.P3(Con) | Remove from Use? Q11.P2.P3(Con) | Standard Acceptance Criteria Remove from Use? Q11.P2.P3(Con) Equal to No
Hemocue Lot to Lot | Equipment Problem Log P2.P3(Con) | Service Called? Q14.P2.P3(Con) | Standard Acceptance Criteria Service Called? Q14.P2.P3(Con) Equal to Yes
Hemocue Lot to Lot | Equipment Problem Log P2.P3(Con) | Date Service Called Q15.P2.P3(Con) | Standard Acceptance Criteria Date Service Called Q15.P2.P3(Con) Less than or equal to Today
Hemocue Lot to Lot | Equipment Problem Log P2.P3(Con) | Equipment Removed? Q18.P2.P3(Con) | Standard Acceptance Criteria Equipment Removed? Q18.P2.P3(Con) Equal to No
Hemocue Lot to Lot | Equipment Problem Log P2.P3(Con) | Date Removed Q19.P2.P3(Con) | Standard Acceptance Criteria Date Removed Q19.P2.P3(Con) Less than or equal to Today
Hemocue Lot to Lot | Equipment Problem Log P2.P3(Con) | Date Returned Q20.P2.P3(Con) | Standard Acceptance Criteria Date Returned Q20.P2.P3(Con) Less than or equal to Today
Hemocue Lot to Lot | Equipment Problem Log P2.P3(Con) | Qualification Validation Required? Q21.P2.P3(Con) | Standard Acceptance Criteria Qualification Validation Required? Q21.P2.P3(Con) Equal to Yes
Hemocue Lot to Lot | Equipment Problem Log P2.P3(Con) | Date Qualification Completed Q22.P2.P3(Con) | Standard Acceptance Criteria Date Qualification Completed Q22.P2.P3(Con) Less than or equal to Today
Hemocue Lot to Lot | Equipment Problem Log P2.P3(Con) | Qualification/Validation Notes Q23.P2.P3(Con) | Standard Acceptance Criteria Qualification/Validation Notes Q23.P2.P3(Con) undefined undefined
Hemocue Lot to Lot | Equipment Problem Log P2.P3(Con) | Qualification/Validation Acceptable? Q24.P2.P3(Con) | Standard Acceptance Criteria Qualification/Validation Acceptable? Q24.P2.P3(Con) Equal to Yes
Hemocue Lot to Lot | Equipment Problem Log P2.P3(Con) | Date Approved Fo Use Q25.P2.P3(Con) | Standard Acceptance Criteria Date Approved Fo Use Q25.P2.P3(Con) Less than or equal to Today
Hemocue Lot to Lot | Problem with QC P3 | Problems with QC Q1.P3 | Standard Acceptance Criteria Problems with QC Q1.P3 Equal to No
Hemocue Lot to Lot | Equipment Problem Log P3.P2(Con) | Date of Problem Q1.P3.P2(Con) | Standard Acceptance Criteria Date of Problem Q1.P3.P2(Con) Less than or equal to Today
Hemocue Lot to Lot | Equipment Problem Log P3.P2(Con) | Date of Trouble Shooting Q8.P3.P2(Con) | Standard Acceptance Criteria Date of Trouble Shooting Q8.P3.P2(Con) Less than or equal to Today
Hemocue Lot to Lot | Equipment Problem Log P3.P2(Con) | Problem Resolved? Q10.P3.P2(Con) | Standard Acceptance Criteria Problem Resolved? Q10.P3.P2(Con) Equal to Yes
Hemocue Lot to Lot | Equipment Problem Log P3.P2(Con) | Remove from Use? Q11.P3.P2(Con) | Standard Acceptance Criteria Remove from Use? Q11.P3.P2(Con) Equal to No
Hemocue Lot to Lot | Equipment Problem Log P3.P2(Con) | Service Called? Q14.P3.P2(Con) | Standard Acceptance Criteria Service Called? Q14.P3.P2(Con) Equal to Yes
Hemocue Lot to Lot | Equipment Problem Log P3.P2(Con) | Date Service Called Q15.P3.P2(Con) | Standard Acceptance Criteria Date Service Called Q15.P3.P2(Con) Less than or equal to Today
Hemocue Lot to Lot | Equipment Problem Log P3.P2(Con) | Equipment Removed? Q18.P3.P2(Con) | Standard Acceptance Criteria Equipment Removed? Q18.P3.P2(Con) Equal to No
Hemocue Lot to Lot | Equipment Problem Log P3.P2(Con) | Date Removed Q19.P3.P2(Con) | Standard Acceptance Criteria Date Removed Q19.P3.P2(Con) Less than or equal to Today
Hemocue Lot to Lot | Equipment Problem Log P3.P2(Con) | Date Returned Q20.P3.P2(Con) | Standard Acceptance Criteria Date Returned Q20.P3.P2(Con) Less than or equal to Today
Hemocue Lot to Lot | Equipment Problem Log P3.P2(Con) | Qualification Validation Required? Q21.P3.P2(Con) | Standard Acceptance Criteria Qualification Validation Required? Q21.P3.P2(Con) Equal to Yes
Hemocue Lot to Lot | Equipment Problem Log P3.P2(Con) | Date Qualification Completed Q22.P3.P2(Con) | Standard Acceptance Criteria Date Qualification Completed Q22.P3.P2(Con) Less than or equal to Today
Hemocue Lot to Lot | Equipment Problem Log P3.P2(Con) | Qualification/Validation Notes Q23.P3.P2(Con) | Standard Acceptance Criteria Qualification/Validation Notes Q23.P3.P2(Con) undefined undefined
Hemocue Lot to Lot | Equipment Problem Log P3.P2(Con) | Qualification/Validation Acceptable? Q24.P3.P2(Con) | Standard Acceptance Criteria Qualification/Validation Acceptable? Q24.P3.P2(Con) Equal to Yes
Hemocue Lot to Lot | Equipment Problem Log P3.P2(Con) | Date Approved Fo Use Q25.P3.P2(Con) | Standard Acceptance Criteria Date Approved Fo Use Q25.P3.P2(Con) Less than or equal to Today
Irradiator Daily-inactive | Radsure Information P1 | Radure Indicator Expiration Date Q2.P1 | Standard Acceptance Criteria Radure Indicator Expiration Date Q2.P1 Greater than or equal to Today
Irradiator Daily-inactive | Radsure Information P1 | Lamp Test Completed Q3.P1 | Standard Acceptance Criteria Lamp Test Completed Q3.P1 Equal to Performed - OK
Compomat G5 Daily | Compomat G5 Daily P1 | Start up G5 Q1.P1 | Standard Acceptance Criteria Start up G5 Q1.P1 Equal to Passed
Compomat G5 Daily | Compomat G5 Daily P1 | Wipe down G5 Q2.P1 | Standard Acceptance Criteria Wipe down G5 Q2.P1 Equal to Performed
Compomat G5 Daily | Compomat G5 Daily P1 | Calibrate Scales (500 G) Q3.P1 | Standard Acceptance Criteria Calibrate Scales (500 G) Q3.P1 Equal to Performed
Compomat G5 Daily | Compomat G5 Daily P1 | Problem with Equipment Q4.P1 | Standard Acceptance Criteria Problem with Equipment Q4.P1 Equal to No
DO Electronic Balance Daily-inactive | Electronic Balance Daily 300 g P1.P3(Con) | Electronic Balance Daily 300 G Q2.P1.P3(Con) | Standard Acceptance Criteria Electronic Balance Daily 300 G Q2.P1.P3(Con) Between 299.9 to 300.1
DO Electronic Balance Daily-inactive | Electronic Balance Daily 300 g P1.P3(Con) | Problems with Equipment/QC Q1.P1.P3(Con) | Standard Acceptance Criteria Problems with Equipment/QC Q1.P1.P3(Con) Equal to No
DO Electronic Balance Daily-inactive | Electronic Balance Daily P1.P2(Con) | Electronic Balance Daily 200 G Q2.P1.P2(Con) | Standard Acceptance Criteria Electronic Balance Daily 200 G Q2.P1.P2(Con) Between 199.9 to 200.1
DO Electronic Balance Daily-inactive | Electronic Balance Daily P1.P2(Con) | Problem with Equipment Q1.P1.P2(Con) | Standard Acceptance Criteria Problem with Equipment Q1.P1.P2(Con) Equal to No
Platelet Sampling Station Decontamination | Platelet Sampling Station Decontamination Daily P2 | Clean and Disinfect Working Surface Areas Where BacT Alert Bottles Are Prepared Using Disinfectant Solution? Q1.P2 | Standard Acceptance Criteria Clean and Disinfect Working Surface Areas Where BacT Alert Bottles Are Prepared Using Disinfectant Solution? Q1.P2 Equal to Performed
Platelet Sampling Station Decontamination | Platelet Sampling Station Decontamination Daily P2 | Problem with Equipment Q2.P2 | Standard Acceptance Criteria Problem with Equipment Q2.P2 Equal to No
Component Manufacturing Lot Tracking | Problem with QC P1 | Problem with QC Q1.P1 | Standard Acceptance Criteria Problem with QC Q1.P1 Equal to No
Irradiator Weekly | Irradiator Weekly P2 | Cleaned/Disinfected Q1.P2 | Standard Acceptance Criteria Cleaned/Disinfected Q1.P2 Equal to Performed
Irradiator Weekly | Irradiator Weekly P2 | Problem with Equipment Q2.P2 | Standard Acceptance Criteria Problem with Equipment Q2.P2 Equal to No
Compomat G5 Weekly | Compomat G5 Weekly P1 | Inspect/Clean Sealing Heads Q1.P1 | Standard Acceptance Criteria Inspect/Clean Sealing Heads Q1.P1 Equal to Performed
Compomat G5 Weekly | Compomat G5 Weekly P1 | Clean Presses and Slide Q2.P1 | Standard Acceptance Criteria Clean Presses and Slide Q2.P1 Equal to Performed
Compomat G5 Weekly | Compomat G5 Weekly P1 | Problem with Equipment Q3.P1 | Standard Acceptance Criteria Problem with Equipment Q3.P1 Equal to No
DO Electronic Balance Weekly-inactive | Electronic Balance Weekly P1 | Cleaned Q1.P1 | Standard Acceptance Criteria Cleaned Q1.P1 Equal to Performed - OK
DO Electronic Balance Weekly-inactive | Electronic Balance Weekly P1 | Problems with Equipment/QC Q2.P1 | Standard Acceptance Criteria Problems with Equipment/QC Q2.P1 Equal to No
Inactive-Duplicate | QC Performed P1 | Was the QC Performed Q1.P1 | Standard Acceptance Criteria Was the QC Performed Q1.P1 Equal to Yes
Inactive-Duplicate | Table Top Sealer Weekly P1.P2(Con) | Clean/Wipe the Heat Sealing Element Q1.P1.P2(Con) | Standard Acceptance Criteria Clean/Wipe the Heat Sealing Element Q1.P1.P2(Con) Equal to Performed - OK
Inactive-Duplicate | Equipment Problem Log P1.P2(Con).P2(Con) | Date of Problem Q1.P1.P2(Con).P2(Con) | Standard Acceptance Criteria Date of Problem Q1.P1.P2(Con).P2(Con) Less than or equal to Today
Inactive-Duplicate | Equipment Problem Log P1.P2(Con).P2(Con) | Date of Trouble Shooting Q8.P1.P2(Con).P2(Con) | Standard Acceptance Criteria Date of Trouble Shooting Q8.P1.P2(Con).P2(Con) Less than or equal to Today
Inactive-Duplicate | Equipment Problem Log P1.P2(Con).P2(Con) | Problem Resolved? Q10.P1.P2(Con).P2(Con) | Standard Acceptance Criteria Problem Resolved? Q10.P1.P2(Con).P2(Con) Equal to Yes
Inactive-Duplicate | Equipment Problem Log P1.P2(Con).P2(Con) | Remove from Use? Q11.P1.P2(Con).P2(Con) | Standard Acceptance Criteria Remove from Use? Q11.P1.P2(Con).P2(Con) Equal to No
Inactive-Duplicate | Equipment Problem Log P1.P2(Con).P2(Con) | Service Called? Q14.P1.P2(Con).P2(Con) | Standard Acceptance Criteria Service Called? Q14.P1.P2(Con).P2(Con) Equal to Yes
Inactive-Duplicate | Equipment Problem Log P1.P2(Con).P2(Con) | Date Service Called Q15.P1.P2(Con).P2(Con) | Standard Acceptance Criteria Date Service Called Q15.P1.P2(Con).P2(Con) Less than or equal to Today
Inactive-Duplicate | Equipment Problem Log P1.P2(Con).P2(Con) | Equipment Removed? Q18.P1.P2(Con).P2(Con) | Standard Acceptance Criteria Equipment Removed? Q18.P1.P2(Con).P2(Con) Equal to No
Inactive-Duplicate | Equipment Problem Log P1.P2(Con).P2(Con) | Date Removed Q19.P1.P2(Con).P2(Con) | Standard Acceptance Criteria Date Removed Q19.P1.P2(Con).P2(Con) Less than or equal to Today
Inactive-Duplicate | Equipment Problem Log P1.P2(Con).P2(Con) | Date Returned Q20.P1.P2(Con).P2(Con) | Standard Acceptance Criteria Date Returned Q20.P1.P2(Con).P2(Con) Less than or equal to Today
Inactive-Duplicate | Equipment Problem Log P1.P2(Con).P2(Con) | Qualification Validation Required? Q21.P1.P2(Con).P2(Con) | Standard Acceptance Criteria Qualification Validation Required? Q21.P1.P2(Con).P2(Con) Equal to Yes
Inactive-Duplicate | Equipment Problem Log P1.P2(Con).P2(Con) | Date Qualification Completed Q22.P1.P2(Con).P2(Con) | Standard Acceptance Criteria Date Qualification Completed Q22.P1.P2(Con).P2(Con) Less than or equal to Today
Inactive-Duplicate | Equipment Problem Log P1.P2(Con).P2(Con) | Qualification/Validation Notes Q23.P1.P2(Con).P2(Con) | Standard Acceptance Criteria Qualification/Validation Notes Q23.P1.P2(Con).P2(Con) undefined undefined
Inactive-Duplicate | Equipment Problem Log P1.P2(Con).P2(Con) | Qualification/Validation Acceptable? Q24.P1.P2(Con).P2(Con) | Standard Acceptance Criteria Qualification/Validation Acceptable? Q24.P1.P2(Con).P2(Con) Equal to Yes
Inactive-Duplicate | Equipment Problem Log P1.P2(Con).P2(Con) | Date Approved Fo Use Q25.P1.P2(Con).P2(Con) | Standard Acceptance Criteria Date Approved Fo Use Q25.P1.P2(Con).P2(Con) Less than or equal to Today
Inactive-Duplicate | Problem with Equipment P2 | Problems with Equipment/QC Q1.P2 | Standard Acceptance Criteria Problems with Equipment/QC Q1.P2 Equal to No
Inactive-Duplicate | Equipment Problem Log P2.P2(Con) | Date of Problem Q1.P2.P2(Con) | Standard Acceptance Criteria Date of Problem Q1.P2.P2(Con) Less than or equal to Today
Inactive-Duplicate | Equipment Problem Log P2.P2(Con) | Date of Trouble Shooting Q8.P2.P2(Con) | Standard Acceptance Criteria Date of Trouble Shooting Q8.P2.P2(Con) Less than or equal to Today
Inactive-Duplicate | Equipment Problem Log P2.P2(Con) | Problem Resolved? Q10.P2.P2(Con) | Standard Acceptance Criteria Problem Resolved? Q10.P2.P2(Con) Equal to Yes
Inactive-Duplicate | Equipment Problem Log P2.P2(Con) | Remove from Use? Q11.P2.P2(Con) | Standard Acceptance Criteria Remove from Use? Q11.P2.P2(Con) Equal to No
Inactive-Duplicate | Equipment Problem Log P2.P2(Con) | Service Called? Q14.P2.P2(Con) | Standard Acceptance Criteria Service Called? Q14.P2.P2(Con) Equal to Yes
Inactive-Duplicate | Equipment Problem Log P2.P2(Con) | Date Service Called Q15.P2.P2(Con) | Standard Acceptance Criteria Date Service Called Q15.P2.P2(Con) Less than or equal to Today
Inactive-Duplicate | Equipment Problem Log P2.P2(Con) | Equipment Removed? Q18.P2.P2(Con) | Standard Acceptance Criteria Equipment Removed? Q18.P2.P2(Con) Equal to No
Inactive-Duplicate | Equipment Problem Log P2.P2(Con) | Date Removed Q19.P2.P2(Con) | Standard Acceptance Criteria Date Removed Q19.P2.P2(Con) Less than or equal to Today
Inactive-Duplicate | Equipment Problem Log P2.P2(Con) | Date Returned Q20.P2.P2(Con) | Standard Acceptance Criteria Date Returned Q20.P2.P2(Con) Less than or equal to Today
Inactive-Duplicate | Equipment Problem Log P2.P2(Con) | Qualification Validation Required? Q21.P2.P2(Con) | Standard Acceptance Criteria Qualification Validation Required? Q21.P2.P2(Con) Equal to Yes
Inactive-Duplicate | Equipment Problem Log P2.P2(Con) | Date Qualification Completed Q22.P2.P2(Con) | Standard Acceptance Criteria Date Qualification Completed Q22.P2.P2(Con) Less than or equal to Today
Inactive-Duplicate | Equipment Problem Log P2.P2(Con) | Qualification/Validation Notes Q23.P2.P2(Con) | Standard Acceptance Criteria Qualification/Validation Notes Q23.P2.P2(Con) undefined undefined
Inactive-Duplicate | Equipment Problem Log P2.P2(Con) | Qualification/Validation Acceptable? Q24.P2.P2(Con) | Standard Acceptance Criteria Qualification/Validation Acceptable? Q24.P2.P2(Con) Equal to Yes
Inactive-Duplicate | Equipment Problem Log P2.P2(Con) | Date Approved Fo Use Q25.P2.P2(Con) | Standard Acceptance Criteria Date Approved Fo Use Q25.P2.P2(Con) Less than or equal to Today
Freezer Check Weekly | Freezer Check Weekly P1 | Brushed/Defrosted Freezer Q1.P1 | Standard Acceptance Criteria Brushed/Defrosted Freezer Q1.P1 Equal to Performed
Freezer Check Weekly | Freezer Check Weekly P1 | Problem with Equipment Q2.P1 | Standard Acceptance Criteria Problem with Equipment Q2.P1 Equal to No
Amicus PM Annual | PM P1 | PM Performed Q3.P1 | Standard Acceptance Criteria PM Performed Q3.P1 Equal to Performed - OK
Amicus PM Annual | PM P1 | Problem with Equipment Q4.P1 | Standard Acceptance Criteria Problem with Equipment Q4.P1 Equal to No
Sorvall ST40 Centrifuge Bi-Annual | Sorvall ST40 Biannual QC P1 | Clean Ventilation Hole Q1.P1 | Standard Acceptance Criteria Clean Ventilation Hole Q1.P1 Equal to Performed - OK
Sorvall ST40 Centrifuge Bi-Annual | Sorvall ST40 Biannual QC P1 | Check Lid and Safety System Q2.P1 | Standard Acceptance Criteria Check Lid and Safety System Q2.P1 Equal to Performed - OK
Sorvall ST40 Centrifuge Bi-Annual | Sorvall ST40 Biannual QC P1 | Check Electronic Balance Recognition Q3.P1 | Standard Acceptance Criteria Check Electronic Balance Recognition Q3.P1 Equal to Performed - OK
Sorvall ST40 Centrifuge Bi-Annual | Sorvall ST40 Biannual QC P1 | Check Rotor Fixation and Driver Shaft Q4.P1 | Standard Acceptance Criteria Check Rotor Fixation and Driver Shaft Q4.P1 Equal to Performed - OK
Refrigerated Centrifuge As Needed | Refrigerated Centrifuge As Needed P1 | Disinfect the Rotor Q1.P1 | Standard Acceptance Criteria Disinfect the Rotor Q1.P1 Equal to Performed
Refrigerated Centrifuge As Needed | Refrigerated Centrifuge As Needed P1 | Defrost the Rotor Chamber Q2.P1 | Standard Acceptance Criteria Defrost the Rotor Chamber Q2.P1 Equal to Performed
Refrigerated Centrifuge As Needed | Refrigerated Centrifuge As Needed P1 | Wash the Spindle Q3.P1 | Standard Acceptance Criteria Wash the Spindle Q3.P1 Equal to Performed
Refrigerated Centrifuge As Needed | Refrigerated Centrifuge As Needed P1 | Check the Gas Springs Q4.P1 | Standard Acceptance Criteria Check the Gas Springs Q4.P1 Equal to Performed
Refrigerated Centrifuge As Needed | Refrigerated Centrifuge As Needed P1 | Check the Gas Spring Bracket Q5.P1 | Standard Acceptance Criteria Check the Gas Spring Bracket Q5.P1 Equal to Performed
Refrigerated Centrifuge As Needed | Refrigerated Centrifuge As Needed P1 | Problem with Equipment Q6.P1 | Standard Acceptance Criteria Problem with Equipment Q6.P1 Equal to No
RPM Calibration Quarterly | Quarterly RPM Calibration P1.P2(Con) | Tachometer RPM Reading Q1.P1.P2(Con) | Standard Acceptance Criteria Tachometer RPM Reading Q1.P1.P2(Con) Between 3400 to 3599
RPM Calibration Quarterly | Quarterly RPM Calibration P1.P3(Con) | Tachometer RPM Reading Q1.P1.P3(Con) | Standard Acceptance Criteria Tachometer RPM Reading Q1.P1.P3(Con) Between 3700 to 3900
Centrifuge Timer Calibration Quarterly | Quarterly Timer Calibration (Serofuge, Cell Washer) P1.P2(Con) | Stopwatch Timer Reading at 10 Seconds Q1.P1.P2(Con) | Standard Acceptance Criteria Stopwatch Timer Reading at 10 Seconds Q1.P1.P2(Con) Between 9 to 11
Centrifuge Timer Calibration Quarterly | Quarterly Timer Calibration (Serofuge, Cell Washer) P1.P2(Con) | Stopwatch Timer Reading at 15 Seconds Q2.P1.P2(Con) | Standard Acceptance Criteria Stopwatch Timer Reading at 15 Seconds Q2.P1.P2(Con) Between 14 to 16
Centrifuge Timer Calibration Quarterly | Quarterly Timer Calibration (Serofuge, Cell Washer) P1.P2(Con) | Stopwatch Timer Reading at 20 Seconds Q3.P1.P2(Con) | Standard Acceptance Criteria Stopwatch Timer Reading at 20 Seconds Q3.P1.P2(Con) Between 19 to 21
Centrifuge Timer Calibration Quarterly | Quarterly Timer Calibration (Serofuge, Cell Washer) P1.P2(Con) | Stopwatch Timer Reading at 30 Seconds Q4.P1.P2(Con) | Standard Acceptance Criteria Stopwatch Timer Reading at 30 Seconds Q4.P1.P2(Con) Between 29 to 31
Centrifuge Timer Calibration Quarterly | Quarterly Timer Calibration (ST 40, +RC3BP) P1.P2(Con) | Stopwatch Timer Reading at 5 Minutes Q1.P1.P2(Con) | Standard Acceptance Criteria Stopwatch Timer Reading at 5 Minutes Q1.P1.P2(Con) Between 4.5 to 5.5
Lot Release Checklist Daily | Daily Lot Release Checklist P1 | Review and Correct All In-process Visits in EDD Q1.P1 | Standard Acceptance Criteria Review and Correct All In-process Visits in EDD Q1.P1 Equal to Performed - OK
Lot Release Checklist Daily | Daily Lot Release Checklist P1 | Review and Correct All In-process Donor Draws in EDD Q2.P1 | Standard Acceptance Criteria Review and Correct All In-process Donor Draws in EDD Q2.P1 Equal to Performed - OK
Lot Release Checklist Daily | Daily Lot Release Checklist P1 | Review and Correct Any Incorrectly Formatted Accession Numbers in EDD Q3.P1 | Standard Acceptance Criteria Review and Correct Any Incorrectly Formatted Accession Numbers in EDD Q3.P1 Equal to Performed - OK
Lot Release Checklist Daily | Daily Lot Release Checklist P1 | Review Donors Deferred for Incomplete Questionnaires and Ensure No Product Was Collected Q4.P1 | Standard Acceptance Criteria Review Donors Deferred for Incomplete Questionnaires and Ensure No Product Was Collected Q4.P1 Equal to Performed - OK
Lot Release Checklist Daily | Daily Lot Release Checklist P1 | Review Deferred Donors and Correct Any Future Deferrals in EDD Q5.P1 | Standard Acceptance Criteria Review Deferred Donors and Correct Any Future Deferrals in EDD Q5.P1 Equal to Performed - OK
Lot Release Checklist Daily | Daily Lot Release Checklist P1 | Print Donor Reactions and Provide to TMP for Review Q6.P1 | Standard Acceptance Criteria Print Donor Reactions and Provide to TMP for Review Q6.P1 Equal to Performed - OK
Lot Release Checklist Daily | Daily Lot Release Checklist P1 | Review and Correct All Duplicate Donors in EDD Q7.P1 | Standard Acceptance Criteria Review and Correct All Duplicate Donors in EDD Q7.P1 Equal to Performed - OK
Lot Release Checklist Daily | Daily Lot Release Checklist P1 | Review and Complete All Incomplete Interviews in DD Q8.P1 | Standard Acceptance Criteria Review and Complete All Incomplete Interviews in DD Q8.P1 Equal to Performed - OK
Lot Release Checklist Daily | Daily Lot Release Checklist P1 | Review Unacceptable Donor Physical Results in EDD Q9.P1 | Standard Acceptance Criteria Review Unacceptable Donor Physical Results in EDD Q9.P1 Equal to Performed - OK
Lot Release Checklist Daily | Daily Lot Release Checklist P1 | Review Sigma QC Records Q10.P1 | Standard Acceptance Criteria Review Sigma QC Records Q10.P1 Equal to Performed - OK
Lot Release Checklist Daily | Daily Lot Release Checklist P1 | Review Rees Report Q11.P1 | Standard Acceptance Criteria Review Rees Report Q11.P1 Equal to Performed - OK
Lot Release Checklist Daily | Daily Lot Release Checklist P1 | Complete Platelet Bag Counts and Notify CM of Any Products to Be Split Q12.P1 | Standard Acceptance Criteria Complete Platelet Bag Counts and Notify CM of Any Products to Be Split Q12.P1 Equal to Performed - OK
Lot Release Checklist Daily | Daily Lot Release Checklist P1 | Add a PLT7D Deferral in Donor Profile If PLTs Is Split. Q13.P1 | Standard Acceptance Criteria Add a PLT7D Deferral in Donor Profile If PLTs Is Split. Q13.P1 Equal to Performed - OK
Lot Release Checklist Daily | Daily Lot Release Checklist P1 | Follow up on Components Quarantine Shelves (Monday-Friday) Q14.P1 | Standard Acceptance Criteria Follow up on Components Quarantine Shelves (Monday-Friday) Q14.P1 Equal to Performed - OK
Lot Release Checklist Daily | Any components requiring follow up? P2 | Any Components Requiring Follow Up? Q1.P2 | Standard Acceptance Criteria Any Components Requiring Follow Up? Q1.P2 Equal to No
Amicus Monthly | Amicus Monthly P1 | Pumps Cleaned Q1.P1 | Standard Acceptance Criteria Pumps Cleaned Q1.P1 Equal to Performed - OK
Amicus Monthly | Amicus Monthly P1 | Pump Head Recess Cleaned Q2.P1 | Standard Acceptance Criteria Pump Head Recess Cleaned Q2.P1 Equal to Yes
Amicus Monthly | Amicus Monthly P1 | Clamps Cleaned Q3.P1 | Standard Acceptance Criteria Clamps Cleaned Q3.P1 Equal to Performed - OK
Amicus Monthly | Amicus Monthly P1 | Air Detector Cleaned Q4.P1 | Standard Acceptance Criteria Air Detector Cleaned Q4.P1 Equal to Yes
Amicus Monthly | Amicus Monthly P1 | Optical Sensor Cleaned Q5.P1 | Standard Acceptance Criteria Optical Sensor Cleaned Q5.P1 Equal to Yes
Amicus Monthly | Amicus Monthly P1 | Cassette Plate Cleaned Q6.P1 | Standard Acceptance Criteria Cassette Plate Cleaned Q6.P1 Equal to Yes
Amicus Monthly | Amicus Monthly P1 | Problem with Equipment Q7.P1 | Standard Acceptance Criteria Problem with Equipment Q7.P1 Equal to No
Cooler QC 1 RBC Annual | Annual Cooler Quality Control 1 RBC P1 | Recorded Temperature of Unit Q2.P1 | Standard Acceptance Criteria Recorded Temperature of Unit Q2.P1 Between 1 to 6
Cooler QC 1 RBC Annual | Annual Cooler Quality Control 1 RBC P1 | Recorded Temperature of Unit Q20.P1 | Standard Acceptance Criteria Recorded Temperature of Unit Q20.P1 Between 1 to 6
Cooler QC 1 RBC Annual | Annual Cooler Quality Control 1 RBC P1 | Recorded Temperature of Unit Q4.P1 | Standard Acceptance Criteria Recorded Temperature of Unit Q4.P1 Between 1 to 6
Cooler QC 1 RBC Annual | Annual Cooler Quality Control 1 RBC P1 | Recorded Temperature of Unit Q6.P1 | Standard Acceptance Criteria Recorded Temperature of Unit Q6.P1 Between 1 to 6
Cooler QC 1 RBC Annual | Annual Cooler Quality Control 1 RBC P1 | Recorded Temperature of Unit Q8.P1 | Standard Acceptance Criteria Recorded Temperature of Unit Q8.P1 Between 1 to 6
Cooler QC 1 RBC Annual | Annual Cooler Quality Control 1 RBC P1 | Recorded Temperature of Unit Q10.P1 | Standard Acceptance Criteria Recorded Temperature of Unit Q10.P1 Between 1 to 6
Cooler QC 1 RBC Annual | Annual Cooler Quality Control 1 RBC P1 | Recorded Temperature of Unit Q12.P1 | Standard Acceptance Criteria Recorded Temperature of Unit Q12.P1 Between 1 to 6
Cooler QC 1 RBC Annual | Annual Cooler Quality Control 1 RBC P1 | Recorded Temperature of Unit Q14.P1 | Standard Acceptance Criteria Recorded Temperature of Unit Q14.P1 Between 1 to 6
Cooler QC 1 RBC Annual | Annual Cooler Quality Control 1 RBC P1 | Recorded Temperature of Unit Q16.P1 | Standard Acceptance Criteria Recorded Temperature of Unit Q16.P1 Between 1 to 6
Cooler QC 1 RBC Annual | Annual Cooler Quality Control 1 RBC P1 | Recorded Temperature of Unit Q18.P1 | Standard Acceptance Criteria Recorded Temperature of Unit Q18.P1 Between 1 to 6
Cooler QC 1 Plasma Annual | Annual Cooler Quality Control 1 Plasma P1 | Recorded Temperature of Unit Q2.P1 | Standard Acceptance Criteria Recorded Temperature of Unit Q2.P1 Between 1 to 6
Cooler QC 1 Plasma Annual | Annual Cooler Quality Control 1 Plasma P1 | Recorded Temperature of Unit Q20.P1 | Standard Acceptance Criteria Recorded Temperature of Unit Q20.P1 Between 1 to 6
Cooler QC 1 Plasma Annual | Annual Cooler Quality Control 1 Plasma P1 | Recorded Temperature of Unit Q4.P1 | Standard Acceptance Criteria Recorded Temperature of Unit Q4.P1 Between 1 to 6
Cooler QC 1 Plasma Annual | Annual Cooler Quality Control 1 Plasma P1 | Recorded Temperature of Unit Q6.P1 | Standard Acceptance Criteria Recorded Temperature of Unit Q6.P1 Between 1 to 6
Cooler QC 1 Plasma Annual | Annual Cooler Quality Control 1 Plasma P1 | Recorded Temperature of Unit Q8.P1 | Standard Acceptance Criteria Recorded Temperature of Unit Q8.P1 Between 1 to 6
Cooler QC 1 Plasma Annual | Annual Cooler Quality Control 1 Plasma P1 | Recorded Temperature of Unit Q10.P1 | Standard Acceptance Criteria Recorded Temperature of Unit Q10.P1 Between 1 to 6
Cooler QC 1 Plasma Annual | Annual Cooler Quality Control 1 Plasma P1 | Recorded Temperature of Unit Q12.P1 | Standard Acceptance Criteria Recorded Temperature of Unit Q12.P1 Between 1 to 6
Cooler QC 1 Plasma Annual | Annual Cooler Quality Control 1 Plasma P1 | Recorded Temperature of Unit Q14.P1 | Standard Acceptance Criteria Recorded Temperature of Unit Q14.P1 Between 1 to 6
Cooler QC 1 Plasma Annual | Annual Cooler Quality Control 1 Plasma P1 | Recorded Temperature of Unit Q16.P1 | Standard Acceptance Criteria Recorded Temperature of Unit Q16.P1 Between 1 to 6
Cooler QC 1 Plasma Annual | Annual Cooler Quality Control 1 Plasma P1 | Recorded Temperature of Unit Q18.P1 | Standard Acceptance Criteria Recorded Temperature of Unit Q18.P1 Between 1 to 6
Cooler QC 3 RBC/3 Plasma Annual | Annual Cooler Quality Control 3 RBC/3 Plasma P1 | Recorded Temperature of Unit Q2.P1 | Standard Acceptance Criteria Recorded Temperature of Unit Q2.P1 Between 1 to 6
Cooler QC 3 RBC/3 Plasma Annual | Annual Cooler Quality Control 3 RBC/3 Plasma P1 | Recorded Temperature of Unit Q20.P1 | Standard Acceptance Criteria Recorded Temperature of Unit Q20.P1 Between 1 to 6
Cooler QC 3 RBC/3 Plasma Annual | Annual Cooler Quality Control 3 RBC/3 Plasma P1 | Recorded Temperature of Unit Q4.P1 | Standard Acceptance Criteria Recorded Temperature of Unit Q4.P1 Between 1 to 6
Cooler QC 3 RBC/3 Plasma Annual | Annual Cooler Quality Control 3 RBC/3 Plasma P1 | Recorded Temperature of Unit Q6.P1 | Standard Acceptance Criteria Recorded Temperature of Unit Q6.P1 Between 1 to 6
Cooler QC 3 RBC/3 Plasma Annual | Annual Cooler Quality Control 3 RBC/3 Plasma P1 | Recorded Temperature of Unit Q8.P1 | Standard Acceptance Criteria Recorded Temperature of Unit Q8.P1 Between 1 to 6
Cooler QC 3 RBC/3 Plasma Annual | Annual Cooler Quality Control 3 RBC/3 Plasma P1 | Recorded Temperature of Unit Q10.P1 | Standard Acceptance Criteria Recorded Temperature of Unit Q10.P1 Between 1 to 6
Cooler QC 3 RBC/3 Plasma Annual | Annual Cooler Quality Control 3 RBC/3 Plasma P1 | Recorded Temperature of Unit Q12.P1 | Standard Acceptance Criteria Recorded Temperature of Unit Q12.P1 Between 1 to 6
Cooler QC 3 RBC/3 Plasma Annual | Annual Cooler Quality Control 3 RBC/3 Plasma P1 | Recorded Temperature of Unit Q14.P1 | Standard Acceptance Criteria Recorded Temperature of Unit Q14.P1 Between 1 to 6
Cooler QC 3 RBC/3 Plasma Annual | Annual Cooler Quality Control 3 RBC/3 Plasma P1 | Recorded Temperature of Unit Q16.P1 | Standard Acceptance Criteria Recorded Temperature of Unit Q16.P1 Between 1 to 6
Cooler QC 3 RBC/3 Plasma Annual | Annual Cooler Quality Control 3 RBC/3 Plasma P1 | Recorded Temperature of Unit Q18.P1 | Standard Acceptance Criteria Recorded Temperature of Unit Q18.P1 Between 1 to 6
Cooler QC 1 Sim Plt Pool/Plt Apheresis Annual | Annual Cooler QC - 1 Sim Plt Pool or Plt Apheresis P1 | Recorded Temperature of Unit Q2.P1 | Standard Acceptance Criteria Recorded Temperature of Unit Q2.P1 Between 20 to 24
Cooler QC 1 Sim Plt Pool/Plt Apheresis Annual | Annual Cooler QC - 1 Sim Plt Pool or Plt Apheresis P1 | Recorded Temperature of Unit Q20.P1 | Standard Acceptance Criteria Recorded Temperature of Unit Q20.P1 Between 20 to 24
Cooler QC 1 Sim Plt Pool/Plt Apheresis Annual | Annual Cooler QC - 1 Sim Plt Pool or Plt Apheresis P1 | Recorded Temperature of Unit Q4.P1 | Standard Acceptance Criteria Recorded Temperature of Unit Q4.P1 Between 20 to 24
Cooler QC 1 Sim Plt Pool/Plt Apheresis Annual | Annual Cooler QC - 1 Sim Plt Pool or Plt Apheresis P1 | Recorded Temperature of Unit Q6.P1 | Standard Acceptance Criteria Recorded Temperature of Unit Q6.P1 Between 20 to 24
Cooler QC 1 Sim Plt Pool/Plt Apheresis Annual | Annual Cooler QC - 1 Sim Plt Pool or Plt Apheresis P1 | Recorded Temperature of Unit Q8.P1 | Standard Acceptance Criteria Recorded Temperature of Unit Q8.P1 Between 20 to 24
Cooler QC 1 Sim Plt Pool/Plt Apheresis Annual | Annual Cooler QC - 1 Sim Plt Pool or Plt Apheresis P1 | Recorded Temperature of Unit Q10.P1 | Standard Acceptance Criteria Recorded Temperature of Unit Q10.P1 Between 20 to 24
Cooler QC 1 Sim Plt Pool/Plt Apheresis Annual | Annual Cooler QC - 1 Sim Plt Pool or Plt Apheresis P1 | Recorded Temperature of Unit Q12.P1 | Standard Acceptance Criteria Recorded Temperature of Unit Q12.P1 Between 20 to 24
Cooler QC 1 Sim Plt Pool/Plt Apheresis Annual | Annual Cooler QC - 1 Sim Plt Pool or Plt Apheresis P1 | Recorded Temperature of Unit Q14.P1 | Standard Acceptance Criteria Recorded Temperature of Unit Q14.P1 Between 20 to 24
Cooler QC 1 Sim Plt Pool/Plt Apheresis Annual | Annual Cooler QC - 1 Sim Plt Pool or Plt Apheresis P1 | Recorded Temperature of Unit Q16.P1 | Standard Acceptance Criteria Recorded Temperature of Unit Q16.P1 Between 20 to 24
Cooler QC 1 Sim Plt Pool/Plt Apheresis Annual | Annual Cooler QC - 1 Sim Plt Pool or Plt Apheresis P1 | Recorded Temperature of Unit Q18.P1 | Standard Acceptance Criteria Recorded Temperature of Unit Q18.P1 Between 20 to 24
Spectra Optia Monthly | Optia Monthly P1 | Pump Housing Cleaned Q1.P1 | Standard Acceptance Criteria Pump Housing Cleaned Q1.P1 Equal to Performed
Spectra Optia Monthly | Optia Monthly P1 | Pump Rotors Cleaned Q2.P1 | Standard Acceptance Criteria Pump Rotors Cleaned Q2.P1 Equal to Performed
Spectra Optia Monthly | Optia Monthly P1 | Fluid Leak Detector Cleaned Q3.P1 | Standard Acceptance Criteria Fluid Leak Detector Cleaned Q3.P1 Equal to Performed
Spectra Optia Monthly | Optia Monthly P1 | Light Covers Cleaned Q4.P1 | Standard Acceptance Criteria Light Covers Cleaned Q4.P1 Equal to Performed - OK
Spectra Optia Monthly | Optia Monthly P1 | Filler Cleaned Q6.P1 | Standard Acceptance Criteria Filler Cleaned Q6.P1 Equal to Performed - OK
Spectra Optia Monthly | Optia Monthly P1 | Centrifuge Chamber Cleaned Q5.P1 | Standard Acceptance Criteria Centrifuge Chamber Cleaned Q5.P1 Equal to Performed - OK
Spectra Optia Monthly | Optia Monthly P1 | Problem with Equipment Q7.P1 | Standard Acceptance Criteria Problem with Equipment Q7.P1 Equal to No
Irradiator Dosimetry Annual | QC Performed P1 | Was the QC Performed Q1.P1 | Standard Acceptance Criteria Was the QC Performed Q1.P1 Equal to Yes
Irradiator Dosimetry Annual | Irradiator Annual QC P1.P2(Con) | Irradiator Dosimetry/Time Changed Q1.P1.P2(Con) | Standard Acceptance Criteria Irradiator Dosimetry/Time Changed Q1.P1.P2(Con) Equal to Performed
Irradiator Dosimetry Annual | Irradiator Annual QC P1.P2(Con) | Problems with Equipment/QC Q2.P1.P2(Con) | Standard Acceptance Criteria Problems with Equipment/QC Q2.P1.P2(Con) Equal to No
Cooler QC 3 Sim Plt Pool/Plt Apheresis Annual | Annual Cooler QC - 3 Sim Plt Pool or Plt Apheresis P1 | Recorded Temperature of Unit Q2.P1 | Standard Acceptance Criteria Recorded Temperature of Unit Q2.P1 Between 20 to 24
Cooler QC 3 Sim Plt Pool/Plt Apheresis Annual | Annual Cooler QC - 3 Sim Plt Pool or Plt Apheresis P1 | Recorded Temperature of Unit Q20.P1 | Standard Acceptance Criteria Recorded Temperature of Unit Q20.P1 Between 20 to 24
Cooler QC 3 Sim Plt Pool/Plt Apheresis Annual | Annual Cooler QC - 3 Sim Plt Pool or Plt Apheresis P1 | Recorded Temperature of Unit Q4.P1 | Standard Acceptance Criteria Recorded Temperature of Unit Q4.P1 Between 20 to 24
Cooler QC 3 Sim Plt Pool/Plt Apheresis Annual | Annual Cooler QC - 3 Sim Plt Pool or Plt Apheresis P1 | Recorded Temperature of Unit Q6.P1 | Standard Acceptance Criteria Recorded Temperature of Unit Q6.P1 Between 20 to 24
Cooler QC 3 Sim Plt Pool/Plt Apheresis Annual | Annual Cooler QC - 3 Sim Plt Pool or Plt Apheresis P1 | Recorded Temperature of Unit Q8.P1 | Standard Acceptance Criteria Recorded Temperature of Unit Q8.P1 Between 20 to 24
Cooler QC 3 Sim Plt Pool/Plt Apheresis Annual | Annual Cooler QC - 3 Sim Plt Pool or Plt Apheresis P1 | Recorded Temperature of Unit Q10.P1 | Standard Acceptance Criteria Recorded Temperature of Unit Q10.P1 Between 20 to 24
Cooler QC 3 Sim Plt Pool/Plt Apheresis Annual | Annual Cooler QC - 3 Sim Plt Pool or Plt Apheresis P1 | Recorded Temperature of Unit Q12.P1 | Standard Acceptance Criteria Recorded Temperature of Unit Q12.P1 Between 20 to 24
Cooler QC 3 Sim Plt Pool/Plt Apheresis Annual | Annual Cooler QC - 3 Sim Plt Pool or Plt Apheresis P1 | Recorded Temperature of Unit Q14.P1 | Standard Acceptance Criteria Recorded Temperature of Unit Q14.P1 Between 20 to 24
Cooler QC 3 Sim Plt Pool/Plt Apheresis Annual | Annual Cooler QC - 3 Sim Plt Pool or Plt Apheresis P1 | Recorded Temperature of Unit Q16.P1 | Standard Acceptance Criteria Recorded Temperature of Unit Q16.P1 Between 20 to 24
Cooler QC 3 Sim Plt Pool/Plt Apheresis Annual | Annual Cooler QC - 3 Sim Plt Pool or Plt Apheresis P1 | Recorded Temperature of Unit Q18.P1 | Standard Acceptance Criteria Recorded Temperature of Unit Q18.P1 Between 20 to 24
Investigation Log | Investigation Log P1 | Problem Resolved? Q2.P1 | Standard Acceptance Criteria Problem Resolved? Q2.P1 Equal to Yes
Compomat G5 PM Annual | PM P1 | PM Performed Q3.P1 | Standard Acceptance Criteria PM Performed Q3.P1 Equal to Performed - OK
Compomat G5 PM Annual | PM P1 | Problem with Equipment Q4.P1 | Standard Acceptance Criteria Problem with Equipment Q4.P1 Equal to No
Freezer PM Annual | PM P1 | PM Performed Q3.P1 | Standard Acceptance Criteria PM Performed Q3.P1 Equal to Performed - OK
Freezer PM Annual | PM P1 | Problem with Equipment Q4.P1 | Standard Acceptance Criteria Problem with Equipment Q4.P1 Equal to No
Cooler QC Annual | Annual Cooler Quality Control 1 Plasma P1.P2(Con) | Recorded Temperature of Unit Q2.P1.P2(Con) | Standard Acceptance Criteria Recorded Temperature of Unit Q2.P1.P2(Con) Between 1 to 6
Cooler QC Annual | Annual Cooler Quality Control 1 Plasma P1.P2(Con) | Recorded Temperature of Unit Q20.P1.P2(Con) | Standard Acceptance Criteria Recorded Temperature of Unit Q20.P1.P2(Con) Between 1 to 6
Cooler QC Annual | Annual Cooler Quality Control 1 Plasma P1.P2(Con) | Recorded Temperature of Unit Q4.P1.P2(Con) | Standard Acceptance Criteria Recorded Temperature of Unit Q4.P1.P2(Con) Between 1 to 6
Cooler QC Annual | Annual Cooler Quality Control 1 Plasma P1.P2(Con) | Recorded Temperature of Unit Q6.P1.P2(Con) | Standard Acceptance Criteria Recorded Temperature of Unit Q6.P1.P2(Con) Between 1 to 6
Cooler QC Annual | Annual Cooler Quality Control 1 Plasma P1.P2(Con) | Recorded Temperature of Unit Q8.P1.P2(Con) | Standard Acceptance Criteria Recorded Temperature of Unit Q8.P1.P2(Con) Between 1 to 6
Cooler QC Annual | Annual Cooler Quality Control 1 Plasma P1.P2(Con) | Recorded Temperature of Unit Q10.P1.P2(Con) | Standard Acceptance Criteria Recorded Temperature of Unit Q10.P1.P2(Con) Between 1 to 6
Cooler QC Annual | Annual Cooler Quality Control 1 Plasma P1.P2(Con) | Recorded Temperature of Unit Q12.P1.P2(Con) | Standard Acceptance Criteria Recorded Temperature of Unit Q12.P1.P2(Con) Between 1 to 6
Cooler QC Annual | Annual Cooler Quality Control 1 Plasma P1.P2(Con) | Recorded Temperature of Unit Q14.P1.P2(Con) | Standard Acceptance Criteria Recorded Temperature of Unit Q14.P1.P2(Con) Between 1 to 6
Cooler QC Annual | Annual Cooler Quality Control 1 Plasma P1.P2(Con) | Recorded Temperature of Unit Q16.P1.P2(Con) | Standard Acceptance Criteria Recorded Temperature of Unit Q16.P1.P2(Con) Between 1 to 6
Cooler QC Annual | Annual Cooler Quality Control 1 Plasma P1.P2(Con) | Recorded Temperature of Unit Q18.P1.P2(Con) | Standard Acceptance Criteria Recorded Temperature of Unit Q18.P1.P2(Con) Between 1 to 6
Cooler QC Annual | Annual Cooler Quality Control 1 Plasma P1.P2(Con) | Problem with Equipment Q21.P1.P2(Con) | Standard Acceptance Criteria Problem with Equipment Q21.P1.P2(Con) Equal to No
Cooler QC Annual | Annual Cooler Quality Control 1 RBC P1.P3(Con) | Recorded Temperature of Unit Q2.P1.P3(Con) | Standard Acceptance Criteria Recorded Temperature of Unit Q2.P1.P3(Con) Between 1 to 6
Cooler QC Annual | Annual Cooler Quality Control 1 RBC P1.P3(Con) | Recorded Temperature of Unit Q20.P1.P3(Con) | Standard Acceptance Criteria Recorded Temperature of Unit Q20.P1.P3(Con) Between 1 to 6
Cooler QC Annual | Annual Cooler Quality Control 1 RBC P1.P3(Con) | Recorded Temperature of Unit Q4.P1.P3(Con) | Standard Acceptance Criteria Recorded Temperature of Unit Q4.P1.P3(Con) Between 1 to 6
Cooler QC Annual | Annual Cooler Quality Control 1 RBC P1.P3(Con) | Recorded Temperature of Unit Q6.P1.P3(Con) | Standard Acceptance Criteria Recorded Temperature of Unit Q6.P1.P3(Con) Between 1 to 6
Cooler QC Annual | Annual Cooler Quality Control 1 RBC P1.P3(Con) | Recorded Temperature of Unit Q8.P1.P3(Con) | Standard Acceptance Criteria Recorded Temperature of Unit Q8.P1.P3(Con) Between 1 to 6
Cooler QC Annual | Annual Cooler Quality Control 1 RBC P1.P3(Con) | Recorded Temperature of Unit Q10.P1.P3(Con) | Standard Acceptance Criteria Recorded Temperature of Unit Q10.P1.P3(Con) Between 1 to 6
Cooler QC Annual | Annual Cooler Quality Control 1 RBC P1.P3(Con) | Recorded Temperature of Unit Q12.P1.P3(Con) | Standard Acceptance Criteria Recorded Temperature of Unit Q12.P1.P3(Con) Between 1 to 6
Cooler QC Annual | Annual Cooler Quality Control 1 RBC P1.P3(Con) | Recorded Temperature of Unit Q14.P1.P3(Con) | Standard Acceptance Criteria Recorded Temperature of Unit Q14.P1.P3(Con) Between 1 to 6
Cooler QC Annual | Annual Cooler Quality Control 1 RBC P1.P3(Con) | Recorded Temperature of Unit Q16.P1.P3(Con) | Standard Acceptance Criteria Recorded Temperature of Unit Q16.P1.P3(Con) Between 1 to 6
Cooler QC Annual | Annual Cooler Quality Control 1 RBC P1.P3(Con) | Recorded Temperature of Unit Q18.P1.P3(Con) | Standard Acceptance Criteria Recorded Temperature of Unit Q18.P1.P3(Con) Between 1 to 6
Cooler QC Annual | Annual Cooler Quality Control 1 RBC P1.P3(Con) | Problem with Equipment Q21.P1.P3(Con) | Standard Acceptance Criteria Problem with Equipment Q21.P1.P3(Con) Equal to No
Cooler QC Annual | Annual Cooler QC - 1 Sim Plt Pool or Plt Apheresis P1.P4(Con) | Recorded Temperature of Unit Q2.P1.P4(Con) | Standard Acceptance Criteria Recorded Temperature of Unit Q2.P1.P4(Con) Between 20 to 24
Cooler QC Annual | Annual Cooler QC - 1 Sim Plt Pool or Plt Apheresis P1.P4(Con) | Recorded Temperature of Unit Q20.P1.P4(Con) | Standard Acceptance Criteria Recorded Temperature of Unit Q20.P1.P4(Con) Between 20 to 24
Cooler QC Annual | Annual Cooler QC - 1 Sim Plt Pool or Plt Apheresis P1.P4(Con) | Recorded Temperature of Unit Q4.P1.P4(Con) | Standard Acceptance Criteria Recorded Temperature of Unit Q4.P1.P4(Con) Between 20 to 24
Cooler QC Annual | Annual Cooler QC - 1 Sim Plt Pool or Plt Apheresis P1.P4(Con) | Recorded Temperature of Unit Q6.P1.P4(Con) | Standard Acceptance Criteria Recorded Temperature of Unit Q6.P1.P4(Con) Between 20 to 24
Cooler QC Annual | Annual Cooler QC - 1 Sim Plt Pool or Plt Apheresis P1.P4(Con) | Recorded Temperature of Unit Q8.P1.P4(Con) | Standard Acceptance Criteria Recorded Temperature of Unit Q8.P1.P4(Con) Between 20 to 24
Cooler QC Annual | Annual Cooler QC - 1 Sim Plt Pool or Plt Apheresis P1.P4(Con) | Recorded Temperature of Unit Q10.P1.P4(Con) | Standard Acceptance Criteria Recorded Temperature of Unit Q10.P1.P4(Con) Between 20 to 24
Cooler QC Annual | Annual Cooler QC - 1 Sim Plt Pool or Plt Apheresis P1.P4(Con) | Recorded Temperature of Unit Q12.P1.P4(Con) | Standard Acceptance Criteria Recorded Temperature of Unit Q12.P1.P4(Con) Between 20 to 24
Cooler QC Annual | Annual Cooler QC - 1 Sim Plt Pool or Plt Apheresis P1.P4(Con) | Recorded Temperature of Unit Q14.P1.P4(Con) | Standard Acceptance Criteria Recorded Temperature of Unit Q14.P1.P4(Con) Between 20 to 24
Cooler QC Annual | Annual Cooler QC - 1 Sim Plt Pool or Plt Apheresis P1.P4(Con) | Recorded Temperature of Unit Q16.P1.P4(Con) | Standard Acceptance Criteria Recorded Temperature of Unit Q16.P1.P4(Con) Between 20 to 24
Cooler QC Annual | Annual Cooler QC - 1 Sim Plt Pool or Plt Apheresis P1.P4(Con) | Recorded Temperature of Unit Q18.P1.P4(Con) | Standard Acceptance Criteria Recorded Temperature of Unit Q18.P1.P4(Con) Between 20 to 24
Cooler QC Annual | Annual Cooler QC - 1 Sim Plt Pool or Plt Apheresis P1.P4(Con) | Problem with Equipment Q21.P1.P4(Con) | Standard Acceptance Criteria Problem with Equipment Q21.P1.P4(Con) Equal to No
Cooler QC Annual | Annual Cooler Quality Control 3 RBC/3 Plasma P1.P5(Con) | Recorded Temperature of Unit Q2.P1.P5(Con) | Standard Acceptance Criteria Recorded Temperature of Unit Q2.P1.P5(Con) Between 1 to 6
Cooler QC Annual | Annual Cooler Quality Control 3 RBC/3 Plasma P1.P5(Con) | Recorded Temperature of Unit Q20.P1.P5(Con) | Standard Acceptance Criteria Recorded Temperature of Unit Q20.P1.P5(Con) Between 1 to 6
Cooler QC Annual | Annual Cooler Quality Control 3 RBC/3 Plasma P1.P5(Con) | Recorded Temperature of Unit Q4.P1.P5(Con) | Standard Acceptance Criteria Recorded Temperature of Unit Q4.P1.P5(Con) Between 1 to 6
Cooler QC Annual | Annual Cooler Quality Control 3 RBC/3 Plasma P1.P5(Con) | Recorded Temperature of Unit Q6.P1.P5(Con) | Standard Acceptance Criteria Recorded Temperature of Unit Q6.P1.P5(Con) Between 1 to 6
Cooler QC Annual | Annual Cooler Quality Control 3 RBC/3 Plasma P1.P5(Con) | Recorded Temperature of Unit Q8.P1.P5(Con) | Standard Acceptance Criteria Recorded Temperature of Unit Q8.P1.P5(Con) Between 1 to 6
Cooler QC Annual | Annual Cooler Quality Control 3 RBC/3 Plasma P1.P5(Con) | Recorded Temperature of Unit Q10.P1.P5(Con) | Standard Acceptance Criteria Recorded Temperature of Unit Q10.P1.P5(Con) Between 1 to 6
Cooler QC Annual | Annual Cooler Quality Control 3 RBC/3 Plasma P1.P5(Con) | Recorded Temperature of Unit Q12.P1.P5(Con) | Standard Acceptance Criteria Recorded Temperature of Unit Q12.P1.P5(Con) Between 1 to 6
Cooler QC Annual | Annual Cooler Quality Control 3 RBC/3 Plasma P1.P5(Con) | Recorded Temperature of Unit Q14.P1.P5(Con) | Standard Acceptance Criteria Recorded Temperature of Unit Q14.P1.P5(Con) Between 1 to 6
Cooler QC Annual | Annual Cooler Quality Control 3 RBC/3 Plasma P1.P5(Con) | Recorded Temperature of Unit Q16.P1.P5(Con) | Standard Acceptance Criteria Recorded Temperature of Unit Q16.P1.P5(Con) Between 1 to 6
Cooler QC Annual | Annual Cooler Quality Control 3 RBC/3 Plasma P1.P5(Con) | Recorded Temperature of Unit Q18.P1.P5(Con) | Standard Acceptance Criteria Recorded Temperature of Unit Q18.P1.P5(Con) Between 1 to 6
Cooler QC Annual | Annual Cooler Quality Control 3 RBC/3 Plasma P1.P5(Con) | Problem with Equipment Q21.P1.P5(Con) | Standard Acceptance Criteria Problem with Equipment Q21.P1.P5(Con) Equal to No
Cooler QC Annual | Annual Cooler QC - 3 Sim Plt Pool or Plt Apheresis P1.P6(Con) | Recorded Temperature of Unit Q2.P1.P6(Con) | Standard Acceptance Criteria Recorded Temperature of Unit Q2.P1.P6(Con) Between 20 to 24
Cooler QC Annual | Annual Cooler QC - 3 Sim Plt Pool or Plt Apheresis P1.P6(Con) | Recorded Temperature of Unit Q20.P1.P6(Con) | Standard Acceptance Criteria Recorded Temperature of Unit Q20.P1.P6(Con) Between 20 to 24
Cooler QC Annual | Annual Cooler QC - 3 Sim Plt Pool or Plt Apheresis P1.P6(Con) | Recorded Temperature of Unit Q4.P1.P6(Con) | Standard Acceptance Criteria Recorded Temperature of Unit Q4.P1.P6(Con) Between 20 to 24
Cooler QC Annual | Annual Cooler QC - 3 Sim Plt Pool or Plt Apheresis P1.P6(Con) | Recorded Temperature of Unit Q6.P1.P6(Con) | Standard Acceptance Criteria Recorded Temperature of Unit Q6.P1.P6(Con) Between 20 to 24
Cooler QC Annual | Annual Cooler QC - 3 Sim Plt Pool or Plt Apheresis P1.P6(Con) | Recorded Temperature of Unit Q8.P1.P6(Con) | Standard Acceptance Criteria Recorded Temperature of Unit Q8.P1.P6(Con) Between 20 to 24
Cooler QC Annual | Annual Cooler QC - 3 Sim Plt Pool or Plt Apheresis P1.P6(Con) | Recorded Temperature of Unit Q10.P1.P6(Con) | Standard Acceptance Criteria Recorded Temperature of Unit Q10.P1.P6(Con) Between 20 to 24
Cooler QC Annual | Annual Cooler QC - 3 Sim Plt Pool or Plt Apheresis P1.P6(Con) | Recorded Temperature of Unit Q12.P1.P6(Con) | Standard Acceptance Criteria Recorded Temperature of Unit Q12.P1.P6(Con) Between 20 to 24
Cooler QC Annual | Annual Cooler QC - 3 Sim Plt Pool or Plt Apheresis P1.P6(Con) | Recorded Temperature of Unit Q14.P1.P6(Con) | Standard Acceptance Criteria Recorded Temperature of Unit Q14.P1.P6(Con) Between 20 to 24
Cooler QC Annual | Annual Cooler QC - 3 Sim Plt Pool or Plt Apheresis P1.P6(Con) | Recorded Temperature of Unit Q16.P1.P6(Con) | Standard Acceptance Criteria Recorded Temperature of Unit Q16.P1.P6(Con) Between 20 to 24
Cooler QC Annual | Annual Cooler QC - 3 Sim Plt Pool or Plt Apheresis P1.P6(Con) | Recorded Temperature of Unit Q18.P1.P6(Con) | Standard Acceptance Criteria Recorded Temperature of Unit Q18.P1.P6(Con) Between 20 to 24
Cooler QC Annual | Annual Cooler QC - 3 Sim Plt Pool or Plt Apheresis P1.P6(Con) | Problem with Equipment Q21.P1.P6(Con) | Standard Acceptance Criteria Problem with Equipment Q21.P1.P6(Con) Equal to No
Cooler QC Annual | Annual Cooler Quality Control Other P1.P7(Con) | Recorded Temperature of Unit Q2.P1.P7(Con) | Standard Acceptance Criteria Recorded Temperature of Unit Q2.P1.P7(Con) Between 1 to 6
Cooler QC Annual | Annual Cooler Quality Control Other P1.P7(Con) | Recorded Temperature of Unit Q4.P1.P7(Con) | Standard Acceptance Criteria Recorded Temperature of Unit Q4.P1.P7(Con) Between 1 to 6
Cooler QC Annual | Annual Cooler Quality Control Other P1.P7(Con) | Recorded Temperature of Unit Q6.P1.P7(Con) | Standard Acceptance Criteria Recorded Temperature of Unit Q6.P1.P7(Con) Between 1 to 6
Cooler QC Annual | Annual Cooler Quality Control Other P1.P7(Con) | Recorded Temperature of Unit Q8.P1.P7(Con) | Standard Acceptance Criteria Recorded Temperature of Unit Q8.P1.P7(Con) Between 1 to 6
Cooler QC Annual | Annual Cooler Quality Control Other P1.P7(Con) | Recorded Temperature of Unit Q10.P1.P7(Con) | Standard Acceptance Criteria Recorded Temperature of Unit Q10.P1.P7(Con) Between 1 to 6
Cooler QC Annual | Annual Cooler Quality Control Other P1.P7(Con) | Recorded Temperature of Unit Q12.P1.P7(Con) | Standard Acceptance Criteria Recorded Temperature of Unit Q12.P1.P7(Con) Between 1 to 6
Cooler QC Annual | Annual Cooler Quality Control Other P1.P7(Con) | Recorded Temperature of Unit Q14.P1.P7(Con) | Standard Acceptance Criteria Recorded Temperature of Unit Q14.P1.P7(Con) Between 1 to 6
Cooler QC Annual | Annual Cooler Quality Control Other P1.P7(Con) | Recorded Temperature of Unit Q16.P1.P7(Con) | Standard Acceptance Criteria Recorded Temperature of Unit Q16.P1.P7(Con) Between 1 to 6
Cooler QC Annual | Annual Cooler Quality Control Other P1.P7(Con) | Recorded Temperature of Unit Q18.P1.P7(Con) | Standard Acceptance Criteria Recorded Temperature of Unit Q18.P1.P7(Con) Between 1 to 6
Cooler QC Annual | Annual Cooler Quality Control Other P1.P7(Con) | Recorded Temperature of Unit Q20.P1.P7(Con) | Standard Acceptance Criteria Recorded Temperature of Unit Q20.P1.P7(Con) Between 1 to 6
Cooler QC Annual | Annual Cooler Quality Control Other P1.P7(Con) | Problem with Equipment Q22.P1.P7(Con) | Standard Acceptance Criteria Problem with Equipment Q22.P1.P7(Con) Equal to No
Cooler QC Annual | Annual Cooler Quality Control Other P1.P8(Con) | Recorded Temperature of Unit Q2.P1.P8(Con) | Standard Acceptance Criteria Recorded Temperature of Unit Q2.P1.P8(Con) Between 20 to 24
Cooler QC Annual | Annual Cooler Quality Control Other P1.P8(Con) | Recorded Temperature of Unit Q4.P1.P8(Con) | Standard Acceptance Criteria Recorded Temperature of Unit Q4.P1.P8(Con) Between 20 to 24
Cooler QC Annual | Annual Cooler Quality Control Other P1.P8(Con) | Recorded Temperature of Unit Q6.P1.P8(Con) | Standard Acceptance Criteria Recorded Temperature of Unit Q6.P1.P8(Con) Between 20 to 24
Cooler QC Annual | Annual Cooler Quality Control Other P1.P8(Con) | Recorded Temperature of Unit Q8.P1.P8(Con) | Standard Acceptance Criteria Recorded Temperature of Unit Q8.P1.P8(Con) Between 20 to 24
Cooler QC Annual | Annual Cooler Quality Control Other P1.P8(Con) | Recorded Temperature of Unit Q10.P1.P8(Con) | Standard Acceptance Criteria Recorded Temperature of Unit Q10.P1.P8(Con) Between 20 to 24
Cooler QC Annual | Annual Cooler Quality Control Other P1.P8(Con) | Recorded Temperature of Unit Q12.P1.P8(Con) | Standard Acceptance Criteria Recorded Temperature of Unit Q12.P1.P8(Con) Between 20 to 24
Cooler QC Annual | Annual Cooler Quality Control Other P1.P8(Con) | Recorded Temperature of Unit Q14.P1.P8(Con) | Standard Acceptance Criteria Recorded Temperature of Unit Q14.P1.P8(Con) Between 20 to 24
Cooler QC Annual | Annual Cooler Quality Control Other P1.P8(Con) | Recorded Temperature of Unit Q16.P1.P8(Con) | Standard Acceptance Criteria Recorded Temperature of Unit Q16.P1.P8(Con) Between 20 to 24
Cooler QC Annual | Annual Cooler Quality Control Other P1.P8(Con) | Recorded Temperature of Unit Q18.P1.P8(Con) | Standard Acceptance Criteria Recorded Temperature of Unit Q18.P1.P8(Con) Between 20 to 24
Cooler QC Annual | Annual Cooler Quality Control Other P1.P8(Con) | Recorded Temperature of Unit Q20.P1.P8(Con) | Standard Acceptance Criteria Recorded Temperature of Unit Q20.P1.P8(Con) Between 20 to 24
Cooler QC Annual | Annual Cooler Quality Control Other P1.P8(Con) | Problem with Equipment Q22.P1.P8(Con) | Standard Acceptance Criteria Problem with Equipment Q22.P1.P8(Con) Equal to No
Accumet pH Meter As Needed | Accumet pH Meter P1 | Standardized with 4.0 Buffer Q1.P1 | Standard Acceptance Criteria Standardized with 4.0 Buffer Q1.P1 Equal to Performed - OK
Accumet pH Meter As Needed | Accumet pH Meter P1 | Percent Slopes Q2.P1 | Standard Acceptance Criteria Percent Slopes Q2.P1 Between 90 to 102
Accumet pH Meter As Needed | Accumet pH Meter P1 | Standardized with 7.0 Buffer Q3.P1 | Standard Acceptance Criteria Standardized with 7.0 Buffer Q3.P1 Equal to Performed - OK
Accumet pH Meter As Needed | Accumet pH Meter P1 | Percent Slopes Q4.P1 | Standard Acceptance Criteria Percent Slopes Q4.P1 Between 90 to 102
Accumet pH Meter As Needed | Accumet pH Meter P1 | Standardized with 10.0 Buffer Q5.P1 | Standard Acceptance Criteria Standardized with 10.0 Buffer Q5.P1 Equal to Performed - OK
Accumet pH Meter As Needed | Accumet pH Meter P1 | Percent Slopes Q6.P1 | Standard Acceptance Criteria Percent Slopes Q6.P1 Between 90 to 102
Accumet pH Meter As Needed | Accumet pH Meter P1 | Electrode Cleaned Q7.P1 | Standard Acceptance Criteria Electrode Cleaned Q7.P1 Equal to Performed
Accumet pH Meter As Needed | Accumet pH Meter P1 | Problem with Equipment Q8.P1 | Standard Acceptance Criteria Problem with Equipment Q8.P1 Equal to No
Spectra Optia PM Bi-Annual | PM P1 | PM Performed Q3.P1 | Standard Acceptance Criteria PM Performed Q3.P1 Equal to Performed - OK
Spectra Optia PM Bi-Annual | PM P1 | Problem with Equipment Q4.P1 | Standard Acceptance Criteria Problem with Equipment Q4.P1 Equal to No
Trima Seal Safe PM Annual | PM P1 | PM Performed Q3.P1 | Standard Acceptance Criteria PM Performed Q3.P1 Equal to Performed - OK
Trima Seal Safe PM Annual | PM P1 | Problem with Equipment Q4.P1 | Standard Acceptance Criteria Problem with Equipment Q4.P1 Equal to No
Coulter AcT 2 Diff S-Calibration Bi-Annual | Coulter AcT Diff 2 S-Calibration P1 | S-Cal Results? Q1.P1 | Standard Acceptance Criteria S-Cal Results? Q1.P1 Equal to Performed - OK
Coulter AcT 2 Diff S-Calibration Bi-Annual | Coulter AcT Diff 2 S-Calibration P1 | Problem with Equipment Q2.P1 | Standard Acceptance Criteria Problem with Equipment Q2.P1 Equal to No
REES Alarm Activation Review Quarterly | Alarm Activation Review P1 | Overall Review of High and Low Alarm Limits Acceptable Q2.P1 | Standard Acceptance Criteria Overall Review of High and Low Alarm Limits Acceptable Q2.P1 Equal to Acceptable
REES Alarm Activation Review Quarterly | Alarm Activation Review P1 | Problem with Equipment Q4.P1 | Standard Acceptance Criteria Problem with Equipment Q4.P1 Equal to No
NIST Thermometers Calibration Annual | Calibration P1 | Calibration Acceptable Q2.P1 | Standard Acceptance Criteria Calibration Acceptable Q2.P1 Equal to Acceptable
NIST Thermometers Calibration Annual | Calibration P1 | Problem with Equipment Q4.P1 | Standard Acceptance Criteria Problem with Equipment Q4.P1 Equal to No
Teletemp Data Logger Calibration Annual | Calibration P1 | Calibration Acceptable Q2.P1 | Standard Acceptance Criteria Calibration Acceptable Q2.P1 Equal to Acceptable
Teletemp Data Logger Calibration Annual | Calibration P1 | Problem with Equipment Q4.P1 | Standard Acceptance Criteria Problem with Equipment Q4.P1 Equal to No
Timers Calibration Annual | Calibration P1 | Calibration Acceptable Q2.P1 | Standard Acceptance Criteria Calibration Acceptable Q2.P1 Equal to Acceptable
Timers Calibration Annual | Calibration P1 | Problem with Equipment Q4.P1 | Standard Acceptance Criteria Problem with Equipment Q4.P1 Equal to No
Weights Calibration Annual | Calibration P1 | Calibration Acceptable Q2.P1 | Standard Acceptance Criteria Calibration Acceptable Q2.P1 Equal to Acceptable
Weights Calibration Annual | Calibration P1 | Problem with Equipment Q4.P1 | Standard Acceptance Criteria Problem with Equipment Q4.P1 Equal to No
Pipettes Calibration Annual | Calibration P1 | Calibration Acceptable Q2.P1 | Standard Acceptance Criteria Calibration Acceptable Q2.P1 Equal to Acceptable
Pipettes Calibration Annual | Calibration P1 | Problem with Equipment Q4.P1 | Standard Acceptance Criteria Problem with Equipment Q4.P1 Equal to No
Thermometers Calibration Annual | Thermometers Calibration P1 | Temperature Difference Q4.P1 | Standard Acceptance Criteria Temperature Difference Q4.P1 Between -1 to 1
Thermometers Calibration Annual | Thermometers Calibration P1 | Calibration Date Q5.P1 | Standard Acceptance Criteria Calibration Date Q5.P1 Less than or equal to Today
Thermometers Calibration Annual | Thermometers Calibration P1 | Problem with Equipment Q7.P1 | Standard Acceptance Criteria Problem with Equipment Q7.P1 Equal to No
Serofuge Calibration Annual | Annual Serofuge Calibration Test Results P1 | Calibration Acceptable Q4.P1 | Standard Acceptance Criteria Calibration Acceptable Q4.P1 Equal to Acceptable
EP Log (copy) | 11-Re-activate Out of Service Equipment P1.P2(Con).P2(Con).P2(Con) | Re-activate Out of Service Equipment Q1.P1.P2(Con).P2(Con).P2(Con) | Standard Acceptance Criteria Re-activate Out of Service Equipment Q1.P1.P2(Con).P2(Con).P2(Con) Equal to Instructions Acknowledged
EP Log (copy) | 11-Re-activate Out of Service Equipment P1.P2(Con).P2(Con).P3(Con).P2(Con).P2(Con) | Re-activate Out of Service Equipment Q1.P1.P2(Con).P2(Con).P3(Con).P2(Con).P2(Con) | Standard Acceptance Criteria Re-activate Out of Service Equipment Q1.P1.P2(Con).P2(Con).P3(Con).P2(Con).P2(Con) Equal to Instructions Acknowledged
EP Log (copy) | 11-Re-activate Out of Service Equipment P1.P2(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P2(Con) | Re-activate Out of Service Equipment Q1.P1.P2(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P2(Con) | Standard Acceptance Criteria Re-activate Out of Service Equipment Q1.P1.P2(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P2(Con) Equal to Instructions Acknowledged
EP Log (copy) | 11-Re-activate Out of Service Equipment P1.P2(Con).P4(Con).P2(Con).P2(Con).P2(Con).P4(Con).P2(Con) | Re-activate Out of Service Equipment Q1.P1.P2(Con).P4(Con).P2(Con).P2(Con).P2(Con).P4(Con).P2(Con) | Standard Acceptance Criteria Re-activate Out of Service Equipment Q1.P1.P2(Con).P4(Con).P2(Con).P2(Con).P2(Con).P4(Con).P2(Con) Equal to Instructions Acknowledged
EP Log (copy) | 11-Re-activate Out of Service Equipment P1.P2(Con).P4(Con).P2(Con).P2(Con).P2(Con).P4(Con).P3(Con).P2(Con).P2(Con) | Re-activate Out of Service Equipment Q1.P1.P2(Con).P4(Con).P2(Con).P2(Con).P2(Con).P4(Con).P3(Con).P2(Con).P2(Con) | Standard Acceptance Criteria Re-activate Out of Service Equipment Q1.P1.P2(Con).P4(Con).P2(Con).P2(Con).P2(Con).P4(Con).P3(Con).P2(Con).P2(Con) Equal to Instructions Acknowledged
EP Log (copy) | 11-Re-activate Out of Service Equipment P1.P2(Con).P4(Con).P2(Con).P2(Con).P2(Con).P5(Con).P2(Con).P2(Con) | Re-activate Out of Service Equipment Q1.P1.P2(Con).P4(Con).P2(Con).P2(Con).P2(Con).P5(Con).P2(Con).P2(Con) | Standard Acceptance Criteria Re-activate Out of Service Equipment Q1.P1.P2(Con).P4(Con).P2(Con).P2(Con).P2(Con).P5(Con).P2(Con).P2(Con) Equal to Instructions Acknowledged
EP Log (copy) | 11-Re-activate Out of Service Equipment P1.P2(Con).P4(Con).P2(Con).P2(Con).P2(Con).P5(Con).P2(Con).P3(Con).P2(Con).P2(Con) | Re-activate Out of Service Equipment Q1.P1.P2(Con).P4(Con).P2(Con).P2(Con).P2(Con).P5(Con).P2(Con).P3(Con).P2(Con).P2(Con) | Standard Acceptance Criteria Re-activate Out of Service Equipment Q1.P1.P2(Con).P4(Con).P2(Con).P2(Con).P2(Con).P5(Con).P2(Con).P3(Con).P2(Con).P2(Con) Equal to Instructions Acknowledged
EP Log (copy) | 11-Re-activate Out of Service Equipment P1.P2(Con).P4(Con).P2(Con).P2(Con).P2(Con).P5(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con) | Re-activate Out of Service Equipment Q1.P1.P2(Con).P4(Con).P2(Con).P2(Con).P2(Con).P5(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con) | Standard Acceptance Criteria Re-activate Out of Service Equipment Q1.P1.P2(Con).P4(Con).P2(Con).P2(Con).P2(Con).P5(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con) Equal to Instructions Acknowledged
EP Log (copy) | 11-Re-activate Out of Service Equipment P1.P2(Con).P4(Con).P2(Con).P2(Con).P3(Con).P2(Con) | Re-activate Out of Service Equipment Q1.P1.P2(Con).P4(Con).P2(Con).P2(Con).P3(Con).P2(Con) | Standard Acceptance Criteria Re-activate Out of Service Equipment Q1.P1.P2(Con).P4(Con).P2(Con).P2(Con).P3(Con).P2(Con) Equal to Instructions Acknowledged
EP Log (copy) | 11-Re-activate Out of Service Equipment P1.P2(Con).P4(Con).P2(Con).P2(Con).P3(Con).P3(Con).P2(Con).P2(Con) | Re-activate Out of Service Equipment Q1.P1.P2(Con).P4(Con).P2(Con).P2(Con).P3(Con).P3(Con).P2(Con).P2(Con) | Standard Acceptance Criteria Re-activate Out of Service Equipment Q1.P1.P2(Con).P4(Con).P2(Con).P2(Con).P3(Con).P3(Con).P2(Con).P2(Con) Equal to Instructions Acknowledged
EP Log (copy) | 11-Re-activate Out of Service Equipment P1.P2(Con).P4(Con).P3(Con).P2(Con).P2(Con).P2(Con) | Re-activate Out of Service Equipment Q1.P1.P2(Con).P4(Con).P3(Con).P2(Con).P2(Con).P2(Con) | Standard Acceptance Criteria Re-activate Out of Service Equipment Q1.P1.P2(Con).P4(Con).P3(Con).P2(Con).P2(Con).P2(Con) Equal to Instructions Acknowledged
EP Log (copy) | 11-Re-activate Out of Service Equipment P1.P2(Con).P4(Con).P3(Con).P2(Con).P2(Con).P3(Con).P2(Con).P2(Con) | Re-activate Out of Service Equipment Q1.P1.P2(Con).P4(Con).P3(Con).P2(Con).P2(Con).P3(Con).P2(Con).P2(Con) | Standard Acceptance Criteria Re-activate Out of Service Equipment Q1.P1.P2(Con).P4(Con).P3(Con).P2(Con).P2(Con).P3(Con).P2(Con).P2(Con) Equal to Instructions Acknowledged
EP Log (copy) | 11-Re-activate Out of Service Equipment P1.P2(Con).P4(Con).P3(Con).P2(Con).P3(Con).P2(Con).P2(Con) | Re-activate Out of Service Equipment Q1.P1.P2(Con).P4(Con).P3(Con).P2(Con).P3(Con).P2(Con).P2(Con) | Standard Acceptance Criteria Re-activate Out of Service Equipment Q1.P1.P2(Con).P4(Con).P3(Con).P2(Con).P3(Con).P2(Con).P2(Con) Equal to Instructions Acknowledged
EP Log (copy) | 11-Re-activate Out of Service Equipment P1.P2(Con).P4(Con).P3(Con).P3(Con).P2(Con) | Re-activate Out of Service Equipment Q1.P1.P2(Con).P4(Con).P3(Con).P3(Con).P2(Con) | Standard Acceptance Criteria Re-activate Out of Service Equipment Q1.P1.P2(Con).P4(Con).P3(Con).P3(Con).P2(Con) Equal to Instructions Acknowledged
EP Log (copy) | 11-Re-activate Out of Service Equipment P1.P2(Con).P4(Con).P3(Con).P3(Con).P3(Con).P2(Con).P2(Con) | Re-activate Out of Service Equipment Q1.P1.P2(Con).P4(Con).P3(Con).P3(Con).P3(Con).P2(Con).P2(Con) | Standard Acceptance Criteria Re-activate Out of Service Equipment Q1.P1.P2(Con).P4(Con).P3(Con).P3(Con).P3(Con).P2(Con).P2(Con) Equal to Instructions Acknowledged
EP Log (copy) | 11-Re-activate Out of Service Equipment P1.P3(Con).P2(Con).P2(Con).P2(Con).P2(Con).P2(Con) | Re-activate Out of Service Equipment Q1.P1.P3(Con).P2(Con).P2(Con).P2(Con).P2(Con).P2(Con) | Standard Acceptance Criteria Re-activate Out of Service Equipment Q1.P1.P3(Con).P2(Con).P2(Con).P2(Con).P2(Con).P2(Con) Equal to Instructions Acknowledged
EP Log (copy) | 11-Re-activate Out of Service Equipment P1.P3(Con).P2(Con).P2(Con).P2(Con).P2(Con).P3(Con).P2(Con).P2(Con) | Re-activate Out of Service Equipment Q1.P1.P3(Con).P2(Con).P2(Con).P2(Con).P2(Con).P3(Con).P2(Con).P2(Con) | Standard Acceptance Criteria Re-activate Out of Service Equipment Q1.P1.P3(Con).P2(Con).P2(Con).P2(Con).P2(Con).P3(Con).P2(Con).P2(Con) Equal to Instructions Acknowledged
EP Log (copy) | 11-Re-activate Out of Service Equipment P1.P3(Con).P2(Con).P2(Con).P2(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P2(Con) | Re-activate Out of Service Equipment Q1.P1.P3(Con).P2(Con).P2(Con).P2(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P2(Con) | Standard Acceptance Criteria Re-activate Out of Service Equipment Q1.P1.P3(Con).P2(Con).P2(Con).P2(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P2(Con) Equal to Instructions Acknowledged
EP Log (copy) | 11-Re-activate Out of Service Equipment P1.P3(Con).P2(Con).P2(Con).P2(Con).P3(Con).P2(Con).P2(Con) | Re-activate Out of Service Equipment Q1.P1.P3(Con).P2(Con).P2(Con).P2(Con).P3(Con).P2(Con).P2(Con) | Standard Acceptance Criteria Re-activate Out of Service Equipment Q1.P1.P3(Con).P2(Con).P2(Con).P2(Con).P3(Con).P2(Con).P2(Con) Equal to Instructions Acknowledged
EP Log (copy) | 11-Re-activate Out of Service Equipment P1.P3(Con).P2(Con).P2(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P2(Con) | Re-activate Out of Service Equipment Q1.P1.P3(Con).P2(Con).P2(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P2(Con) | Standard Acceptance Criteria Re-activate Out of Service Equipment Q1.P1.P3(Con).P2(Con).P2(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P2(Con) Equal to Instructions Acknowledged
EP Log (copy) | 11-Re-activate Out of Service Equipment P1.P3(Con).P2(Con).P2(Con).P3(Con).P2(Con) | Re-activate Out of Service Equipment Q1.P1.P3(Con).P2(Con).P2(Con).P3(Con).P2(Con) | Standard Acceptance Criteria Re-activate Out of Service Equipment Q1.P1.P3(Con).P2(Con).P2(Con).P3(Con).P2(Con) Equal to Instructions Acknowledged
EP Log (copy) | 11-Re-activate Out of Service Equipment P1.P3(Con).P2(Con).P2(Con).P3(Con).P3(Con).P2(Con).P2(Con) | Re-activate Out of Service Equipment Q1.P1.P3(Con).P2(Con).P2(Con).P3(Con).P3(Con).P2(Con).P2(Con) | Standard Acceptance Criteria Re-activate Out of Service Equipment Q1.P1.P3(Con).P2(Con).P2(Con).P3(Con).P3(Con).P2(Con).P2(Con) Equal to Instructions Acknowledged
EP Log (copy) | 11-Re-activate Out of Service Equipment P1.P3(Con).P2(Con).P2(Con).P3(Con).P3(Con).P2(Con).P3(Con).P2(Con).P2(Con) | Re-activate Out of Service Equipment Q1.P1.P3(Con).P2(Con).P2(Con).P3(Con).P3(Con).P2(Con).P3(Con).P2(Con).P2(Con) | Standard Acceptance Criteria Re-activate Out of Service Equipment Q1.P1.P3(Con).P2(Con).P2(Con).P3(Con).P3(Con).P2(Con).P3(Con).P2(Con).P2(Con) Equal to Instructions Acknowledged
EP Log (copy) | 11-Re-activate Out of Service Equipment P1.P3(Con).P2(Con).P2(Con).P3(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P2(Con) | Re-activate Out of Service Equipment Q1.P1.P3(Con).P2(Con).P2(Con).P3(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P2(Con) | Standard Acceptance Criteria Re-activate Out of Service Equipment Q1.P1.P3(Con).P2(Con).P2(Con).P3(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P2(Con) Equal to Instructions Acknowledged
EP Log (copy) | 11-Re-activate Out of Service Equipment P1.P3(Con).P2(Con).P2(Con).P3(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P2(Con) | Re-activate Out of Service Equipment Q1.P1.P3(Con).P2(Con).P2(Con).P3(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P2(Con) | Standard Acceptance Criteria Re-activate Out of Service Equipment Q1.P1.P3(Con).P2(Con).P2(Con).P3(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P2(Con) Equal to Instructions Acknowledged
EP Log (copy) | 11-Re-activate Out of Service Equipment P1.P3(Con).P2(Con).P2(Con).P3(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P2(Con) | Re-activate Out of Service Equipment Q1.P1.P3(Con).P2(Con).P2(Con).P3(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P2(Con) | Standard Acceptance Criteria Re-activate Out of Service Equipment Q1.P1.P3(Con).P2(Con).P2(Con).P3(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P2(Con) Equal to Instructions Acknowledged
EP Log (copy) | 11-Re-activate Out of Service Equipment P1.P3(Con).P2(Con).P2(Con).P3(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P2(Con) | Re-activate Out of Service Equipment Q1.P1.P3(Con).P2(Con).P2(Con).P3(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P2(Con) | Standard Acceptance Criteria Re-activate Out of Service Equipment Q1.P1.P3(Con).P2(Con).P2(Con).P3(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P2(Con) Equal to Instructions Acknowledged
EP Log (copy) | 11-Re-activate Out of Service Equipment P1.P3(Con).P2(Con).P2(Con).P3(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P2(Con) | Re-activate Out of Service Equipment Q1.P1.P3(Con).P2(Con).P2(Con).P3(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P2(Con) | Standard Acceptance Criteria Re-activate Out of Service Equipment Q1.P1.P3(Con).P2(Con).P2(Con).P3(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P2(Con) Equal to Instructions Acknowledged
EP Log (copy) | 11-Re-activate Out of Service Equipment P1.P3(Con).P2(Con).P2(Con).P3(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con) | Re-activate Out of Service Equipment Q1.P1.P3(Con).P2(Con).P2(Con).P3(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con) | Standard Acceptance Criteria Re-activate Out of Service Equipment Q1.P1.P3(Con).P2(Con).P2(Con).P3(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con) Equal to Instructions Acknowledged
EP Log (copy) | 11-Re-activate Out of Service Equipment P1.P3(Con).P3(Con).P2(Con).P2(Con).P2(Con) | Re-activate Out of Service Equipment Q1.P1.P3(Con).P3(Con).P2(Con).P2(Con).P2(Con) | Standard Acceptance Criteria Re-activate Out of Service Equipment Q1.P1.P3(Con).P3(Con).P2(Con).P2(Con).P2(Con) Equal to Instructions Acknowledged
EP Log (copy) | 11-Re-activate Out of Service Equipment P1.P3(Con).P3(Con).P2(Con).P2(Con).P3(Con).P2(Con).P2(Con) | Re-activate Out of Service Equipment Q1.P1.P3(Con).P3(Con).P2(Con).P2(Con).P3(Con).P2(Con).P2(Con) | Standard Acceptance Criteria Re-activate Out of Service Equipment Q1.P1.P3(Con).P3(Con).P2(Con).P2(Con).P3(Con).P2(Con).P2(Con) Equal to Instructions Acknowledged
EP Log (copy) | 11-Re-activate Out of Service Equipment P1.P3(Con).P3(Con).P2(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P2(Con) | Re-activate Out of Service Equipment Q1.P1.P3(Con).P3(Con).P2(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P2(Con) | Standard Acceptance Criteria Re-activate Out of Service Equipment Q1.P1.P3(Con).P3(Con).P2(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P2(Con) Equal to Instructions Acknowledged
EP Log (copy) | 11-Re-activate Out of Service Equipment P1.P3(Con).P3(Con).P2(Con).P3(Con).P2(Con).P2(Con) | Re-activate Out of Service Equipment Q1.P1.P3(Con).P3(Con).P2(Con).P3(Con).P2(Con).P2(Con) | Standard Acceptance Criteria Re-activate Out of Service Equipment Q1.P1.P3(Con).P3(Con).P2(Con).P3(Con).P2(Con).P2(Con) Equal to Instructions Acknowledged
EP Log (copy) | 11-Re-activate Out of Service Equipment P1.P3(Con).P3(Con).P3(Con).P2(Con) | Re-activate Out of Service Equipment Q1.P1.P3(Con).P3(Con).P3(Con).P2(Con) | Standard Acceptance Criteria Re-activate Out of Service Equipment Q1.P1.P3(Con).P3(Con).P3(Con).P2(Con) Equal to Instructions Acknowledged
EP Log (copy) | 11-Re-activate Out of Service Equipment P1.P3(Con).P3(Con).P3(Con).P3(Con).P2(Con).P2(Con) | Re-activate Out of Service Equipment Q1.P1.P3(Con).P3(Con).P3(Con).P3(Con).P2(Con).P2(Con) | Standard Acceptance Criteria Re-activate Out of Service Equipment Q1.P1.P3(Con).P3(Con).P3(Con).P3(Con).P2(Con).P2(Con) Equal to Instructions Acknowledged
1 - Spectra Cobe Monthly QC | Cobe Monthly P1 | Replace Pressure Pod Seal Q1.P1 | Standard Acceptance Criteria Replace Pressure Pod Seal Q1.P1 Equal to Performed
1 - Spectra Cobe Monthly QC | Cobe Monthly P1 | Front Door Sensors Cleaned Q2.P1 | Standard Acceptance Criteria Front Door Sensors Cleaned Q2.P1 Equal to Performed
1 - Spectra Cobe Monthly QC | Cobe Monthly P1 | Prime Completed Successfully Q3.P1 | Standard Acceptance Criteria Prime Completed Successfully Q3.P1 Equal to Performed
1 - Spectra Cobe Monthly QC | Cobe Monthly P1 | Alarm Check for Access Pressure Sensor Q4.P1 | Standard Acceptance Criteria Alarm Check for Access Pressure Sensor Q4.P1 Equal to Performed
1 - Spectra Cobe Monthly QC | Cobe Monthly P1 | Alarm Check for Return Pressure Sensor Q5.P1 | Standard Acceptance Criteria Alarm Check for Return Pressure Sensor Q5.P1 Equal to Performed
1 - Spectra Cobe Monthly QC | Cobe Monthly P1 | Alarm Fluid Leak Detector Q6.P1 | Standard Acceptance Criteria Alarm Fluid Leak Detector Q6.P1 Equal to Performed
1 - Spectra Cobe Monthly QC | Cobe Monthly P1 | Alarm Check of Return Air Detector Q7.P1 | Standard Acceptance Criteria Alarm Check of Return Air Detector Q7.P1 Equal to Performed
1 - Spectra Cobe Monthly QC | Cobe Monthly P1 | Problem with Equipment Q8.P1 | Standard Acceptance Criteria Problem with Equipment Q8.P1 Equal to No
Hematology Analyzer Bi-Annual Parallels | Hematology Bi-Annual Parallels Instrument 1 P1 | Sample 1 WBC Value Q2.P1 | Standard Acceptance Criteria Sample 1 WBC Value Q2.P1 Between 0 to 999
Hematology Analyzer Bi-Annual Parallels | Hematology Bi-Annual Parallels Instrument 1 P1 | Sample 1 RBC Value Q3.P1 | Standard Acceptance Criteria Sample 1 RBC Value Q3.P1 Between 0 to 999
Hematology Analyzer Bi-Annual Parallels | Hematology Bi-Annual Parallels Instrument 1 P1 | Sample 1 HGB Value Q4.P1 | Standard Acceptance Criteria Sample 1 HGB Value Q4.P1 Between 0 to 999
Hematology Analyzer Bi-Annual Parallels | Hematology Bi-Annual Parallels Instrument 1 P1 | Sample 1 HCT Value Q5.P1 | Standard Acceptance Criteria Sample 1 HCT Value Q5.P1 Between 0 to 999
Hematology Analyzer Bi-Annual Parallels | Hematology Bi-Annual Parallels Instrument 1 P1 | Sample 1 PLT Value Q6.P1 | Standard Acceptance Criteria Sample 1 PLT Value Q6.P1 Between 0 to 999
Hematology Analyzer Bi-Annual Parallels | Hematology Bi-Annual Parallels Instrument 2 P2 | Sample 1 WBC Value Q2.P2 | Standard Acceptance Criteria Sample 1 WBC Value Q2.P2 Between 0 to 999
Hematology Analyzer Bi-Annual Parallels | Hematology Bi-Annual Parallels Instrument 2 P2 | Sample 1 RBC Value Q3.P2 | Standard Acceptance Criteria Sample 1 RBC Value Q3.P2 Between 0 to 999
Hematology Analyzer Bi-Annual Parallels | Hematology Bi-Annual Parallels Instrument 2 P2 | Sample 1 HGB Value Q4.P2 | Standard Acceptance Criteria Sample 1 HGB Value Q4.P2 Between 0 to 999
Hematology Analyzer Bi-Annual Parallels | Hematology Bi-Annual Parallels Instrument 2 P2 | Sample 1 HCT Value Q5.P2 | Standard Acceptance Criteria Sample 1 HCT Value Q5.P2 Between 0 to 999
Hematology Analyzer Bi-Annual Parallels | Hematology Bi-Annual Parallels Instrument 2 P2 | Sample 1 PLT Value Q6.P2 | Standard Acceptance Criteria Sample 1 PLT Value Q6.P2 Between 0 to 999
Hematology Analyzer Bi-Annual Parallels | Hematology Bi-Annual Parallels Instrument 3 P3 | Sample 1 WBC Value Q2.P3 | Standard Acceptance Criteria Sample 1 WBC Value Q2.P3 Between 0 to 999
Hematology Analyzer Bi-Annual Parallels | Hematology Bi-Annual Parallels Instrument 3 P3 | Sample 1 RBC Value Q3.P3 | Standard Acceptance Criteria Sample 1 RBC Value Q3.P3 Between 0 to 999
Hematology Analyzer Bi-Annual Parallels | Hematology Bi-Annual Parallels Instrument 3 P3 | Sample 1 HGB Value Q4.P3 | Standard Acceptance Criteria Sample 1 HGB Value Q4.P3 Between 0 to 999
Hematology Analyzer Bi-Annual Parallels | Hematology Bi-Annual Parallels Instrument 3 P3 | Sample 1 HCT Value Q5.P3 | Standard Acceptance Criteria Sample 1 HCT Value Q5.P3 Between 0 to 999
Hematology Analyzer Bi-Annual Parallels | Hematology Bi-Annual Parallels Instrument 3 P3 | Sample 1 PLT Value Q6.P3 | Standard Acceptance Criteria Sample 1 PLT Value Q6.P3 Between 0 to 999
Hematology Analyzer Bi-Annual Parallels | Hematology Bi-Annual Parallels Instrument 4 P4 | Sample 1 WBC Value Q2.P4 | Standard Acceptance Criteria Sample 1 WBC Value Q2.P4 Between 0 to 999
Hematology Analyzer Bi-Annual Parallels | Hematology Bi-Annual Parallels Instrument 4 P4 | Sample 1 RBC Value Q3.P4 | Standard Acceptance Criteria Sample 1 RBC Value Q3.P4 Between 0 to 999
Hematology Analyzer Bi-Annual Parallels | Hematology Bi-Annual Parallels Instrument 4 P4 | Sample 1 HGB Value Q4.P4 | Standard Acceptance Criteria Sample 1 HGB Value Q4.P4 Between 0 to 999
Hematology Analyzer Bi-Annual Parallels | Hematology Bi-Annual Parallels Instrument 4 P4 | Sample 1 HCT Value Q5.P4 | Standard Acceptance Criteria Sample 1 HCT Value Q5.P4 Between 0 to 999
Hematology Analyzer Bi-Annual Parallels | Hematology Bi-Annual Parallels Instrument 4 P4 | Sample 1 PLT Value Q6.P4 | Standard Acceptance Criteria Sample 1 PLT Value Q6.P4 Between 0 to 999
Hematology Analyzer Bi-Annual Parallels | Hematology Bi-Annual Parallels Instrument 5 P5 | Sample 1 WBC Value Q2.P5 | Standard Acceptance Criteria Sample 1 WBC Value Q2.P5 Between 0 to 999
Hematology Analyzer Bi-Annual Parallels | Hematology Bi-Annual Parallels Instrument 5 P5 | Sample 1 RBC Value Q3.P5 | Standard Acceptance Criteria Sample 1 RBC Value Q3.P5 Between 0 to 999
Hematology Analyzer Bi-Annual Parallels | Hematology Bi-Annual Parallels Instrument 5 P5 | Sample 1 HGB Value Q4.P5 | Standard Acceptance Criteria Sample 1 HGB Value Q4.P5 Between 0 to 999
Hematology Analyzer Bi-Annual Parallels | Hematology Bi-Annual Parallels Instrument 5 P5 | Sample 1 HCT Value Q5.P5 | Standard Acceptance Criteria Sample 1 HCT Value Q5.P5 Between 0 to 999
Hematology Analyzer Bi-Annual Parallels | Hematology Bi-Annual Parallels Instrument 5 P5 | Sample 1 PLT Value Q6.P5 | Standard Acceptance Criteria Sample 1 PLT Value Q6.P5 Between 0 to 999
Equipment Problem Log - Copy RLP | 11-Re-activate Out of Service Equipment P1.P2(Con).P2(Con).P2(Con) | Re-activate Out of Service Equipment Q1.P1.P2(Con).P2(Con).P2(Con) | Standard Acceptance Criteria Re-activate Out of Service Equipment Q1.P1.P2(Con).P2(Con).P2(Con) Equal to Instructions Acknowledged
Equipment Problem Log - Copy RLP | 11-Re-activate Out of Service Equipment P1.P2(Con).P2(Con).P3(Con).P2(Con).P2(Con) | Re-activate Out of Service Equipment Q1.P1.P2(Con).P2(Con).P3(Con).P2(Con).P2(Con) | Standard Acceptance Criteria Re-activate Out of Service Equipment Q1.P1.P2(Con).P2(Con).P3(Con).P2(Con).P2(Con) Equal to Instructions Acknowledged
Equipment Problem Log - Copy RLP | 11-Re-activate Out of Service Equipment P1.P2(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P2(Con) | Re-activate Out of Service Equipment Q1.P1.P2(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P2(Con) | Standard Acceptance Criteria Re-activate Out of Service Equipment Q1.P1.P2(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P2(Con) Equal to Instructions Acknowledged
Equipment Problem Log - Copy RLP | 11-Re-activate Out of Service Equipment P1.P2(Con).P4(Con).P2(Con).P2(Con).P2(Con).P4(Con).P2(Con) | Re-activate Out of Service Equipment Q1.P1.P2(Con).P4(Con).P2(Con).P2(Con).P2(Con).P4(Con).P2(Con) | Standard Acceptance Criteria Re-activate Out of Service Equipment Q1.P1.P2(Con).P4(Con).P2(Con).P2(Con).P2(Con).P4(Con).P2(Con) Equal to Instructions Acknowledged
Equipment Problem Log - Copy RLP | 11-Re-activate Out of Service Equipment P1.P2(Con).P4(Con).P2(Con).P2(Con).P2(Con).P4(Con).P3(Con).P2(Con).P2(Con) | Re-activate Out of Service Equipment Q1.P1.P2(Con).P4(Con).P2(Con).P2(Con).P2(Con).P4(Con).P3(Con).P2(Con).P2(Con) | Standard Acceptance Criteria Re-activate Out of Service Equipment Q1.P1.P2(Con).P4(Con).P2(Con).P2(Con).P2(Con).P4(Con).P3(Con).P2(Con).P2(Con) Equal to Instructions Acknowledged
Equipment Problem Log - Copy RLP | 11-Re-activate Out of Service Equipment P1.P2(Con).P4(Con).P2(Con).P2(Con).P2(Con).P5(Con).P2(Con).P2(Con) | Re-activate Out of Service Equipment Q1.P1.P2(Con).P4(Con).P2(Con).P2(Con).P2(Con).P5(Con).P2(Con).P2(Con) | Standard Acceptance Criteria Re-activate Out of Service Equipment Q1.P1.P2(Con).P4(Con).P2(Con).P2(Con).P2(Con).P5(Con).P2(Con).P2(Con) Equal to Instructions Acknowledged
Equipment Problem Log - Copy RLP | 11-Re-activate Out of Service Equipment P1.P2(Con).P4(Con).P2(Con).P2(Con).P2(Con).P5(Con).P2(Con).P3(Con).P2(Con).P2(Con) | Re-activate Out of Service Equipment Q1.P1.P2(Con).P4(Con).P2(Con).P2(Con).P2(Con).P5(Con).P2(Con).P3(Con).P2(Con).P2(Con) | Standard Acceptance Criteria Re-activate Out of Service Equipment Q1.P1.P2(Con).P4(Con).P2(Con).P2(Con).P2(Con).P5(Con).P2(Con).P3(Con).P2(Con).P2(Con) Equal to Instructions Acknowledged
Equipment Problem Log - Copy RLP | 11-Re-activate Out of Service Equipment P1.P2(Con).P4(Con).P2(Con).P2(Con).P2(Con).P5(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con) | Re-activate Out of Service Equipment Q1.P1.P2(Con).P4(Con).P2(Con).P2(Con).P2(Con).P5(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con) | Standard Acceptance Criteria Re-activate Out of Service Equipment Q1.P1.P2(Con).P4(Con).P2(Con).P2(Con).P2(Con).P5(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con) Equal to Instructions Acknowledged
Equipment Problem Log - Copy RLP | 11-Re-activate Out of Service Equipment P1.P2(Con).P4(Con).P2(Con).P2(Con).P3(Con).P2(Con) | Re-activate Out of Service Equipment Q1.P1.P2(Con).P4(Con).P2(Con).P2(Con).P3(Con).P2(Con) | Standard Acceptance Criteria Re-activate Out of Service Equipment Q1.P1.P2(Con).P4(Con).P2(Con).P2(Con).P3(Con).P2(Con) Equal to Instructions Acknowledged
Equipment Problem Log - Copy RLP | 11-Re-activate Out of Service Equipment P1.P2(Con).P4(Con).P2(Con).P2(Con).P3(Con).P3(Con).P2(Con).P2(Con) | Re-activate Out of Service Equipment Q1.P1.P2(Con).P4(Con).P2(Con).P2(Con).P3(Con).P3(Con).P2(Con).P2(Con) | Standard Acceptance Criteria Re-activate Out of Service Equipment Q1.P1.P2(Con).P4(Con).P2(Con).P2(Con).P3(Con).P3(Con).P2(Con).P2(Con) Equal to Instructions Acknowledged
Equipment Problem Log - Copy RLP | 11-Re-activate Out of Service Equipment P1.P2(Con).P4(Con).P3(Con).P2(Con).P2(Con).P2(Con) | Re-activate Out of Service Equipment Q1.P1.P2(Con).P4(Con).P3(Con).P2(Con).P2(Con).P2(Con) | Standard Acceptance Criteria Re-activate Out of Service Equipment Q1.P1.P2(Con).P4(Con).P3(Con).P2(Con).P2(Con).P2(Con) Equal to Instructions Acknowledged
Equipment Problem Log - Copy RLP | 11-Re-activate Out of Service Equipment P1.P2(Con).P4(Con).P3(Con).P2(Con).P2(Con).P3(Con).P2(Con).P2(Con) | Re-activate Out of Service Equipment Q1.P1.P2(Con).P4(Con).P3(Con).P2(Con).P2(Con).P3(Con).P2(Con).P2(Con) | Standard Acceptance Criteria Re-activate Out of Service Equipment Q1.P1.P2(Con).P4(Con).P3(Con).P2(Con).P2(Con).P3(Con).P2(Con).P2(Con) Equal to Instructions Acknowledged
Equipment Problem Log - Copy RLP | 11-Re-activate Out of Service Equipment P1.P2(Con).P4(Con).P3(Con).P2(Con).P3(Con).P2(Con).P2(Con) | Re-activate Out of Service Equipment Q1.P1.P2(Con).P4(Con).P3(Con).P2(Con).P3(Con).P2(Con).P2(Con) | Standard Acceptance Criteria Re-activate Out of Service Equipment Q1.P1.P2(Con).P4(Con).P3(Con).P2(Con).P3(Con).P2(Con).P2(Con) Equal to Instructions Acknowledged
Equipment Problem Log - Copy RLP | 11-Re-activate Out of Service Equipment P1.P2(Con).P4(Con).P3(Con).P3(Con).P2(Con) | Re-activate Out of Service Equipment Q1.P1.P2(Con).P4(Con).P3(Con).P3(Con).P2(Con) | Standard Acceptance Criteria Re-activate Out of Service Equipment Q1.P1.P2(Con).P4(Con).P3(Con).P3(Con).P2(Con) Equal to Instructions Acknowledged
Equipment Problem Log - Copy RLP | 11-Re-activate Out of Service Equipment P1.P2(Con).P4(Con).P3(Con).P3(Con).P3(Con).P2(Con).P2(Con) | Re-activate Out of Service Equipment Q1.P1.P2(Con).P4(Con).P3(Con).P3(Con).P3(Con).P2(Con).P2(Con) | Standard Acceptance Criteria Re-activate Out of Service Equipment Q1.P1.P2(Con).P4(Con).P3(Con).P3(Con).P3(Con).P2(Con).P2(Con) Equal to Instructions Acknowledged
Equipment Problem Log - Copy RLP | 11-Re-activate Out of Service Equipment P1.P3(Con).P2(Con).P2(Con).P2(Con).P2(Con).P2(Con) | Re-activate Out of Service Equipment Q1.P1.P3(Con).P2(Con).P2(Con).P2(Con).P2(Con).P2(Con) | Standard Acceptance Criteria Re-activate Out of Service Equipment Q1.P1.P3(Con).P2(Con).P2(Con).P2(Con).P2(Con).P2(Con) Equal to Instructions Acknowledged
Equipment Problem Log - Copy RLP | 11-Re-activate Out of Service Equipment P1.P3(Con).P2(Con).P2(Con).P2(Con).P2(Con).P3(Con).P2(Con).P2(Con) | Re-activate Out of Service Equipment Q1.P1.P3(Con).P2(Con).P2(Con).P2(Con).P2(Con).P3(Con).P2(Con).P2(Con) | Standard Acceptance Criteria Re-activate Out of Service Equipment Q1.P1.P3(Con).P2(Con).P2(Con).P2(Con).P2(Con).P3(Con).P2(Con).P2(Con) Equal to Instructions Acknowledged
Equipment Problem Log - Copy RLP | 11-Re-activate Out of Service Equipment P1.P3(Con).P2(Con).P2(Con).P2(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P2(Con) | Re-activate Out of Service Equipment Q1.P1.P3(Con).P2(Con).P2(Con).P2(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P2(Con) | Standard Acceptance Criteria Re-activate Out of Service Equipment Q1.P1.P3(Con).P2(Con).P2(Con).P2(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P2(Con) Equal to Instructions Acknowledged
Equipment Problem Log - Copy RLP | 11-Re-activate Out of Service Equipment P1.P3(Con).P2(Con).P2(Con).P2(Con).P3(Con).P2(Con).P2(Con) | Re-activate Out of Service Equipment Q1.P1.P3(Con).P2(Con).P2(Con).P2(Con).P3(Con).P2(Con).P2(Con) | Standard Acceptance Criteria Re-activate Out of Service Equipment Q1.P1.P3(Con).P2(Con).P2(Con).P2(Con).P3(Con).P2(Con).P2(Con) Equal to Instructions Acknowledged
Equipment Problem Log - Copy RLP | 11-Re-activate Out of Service Equipment P1.P3(Con).P2(Con).P2(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P2(Con) | Re-activate Out of Service Equipment Q1.P1.P3(Con).P2(Con).P2(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P2(Con) | Standard Acceptance Criteria Re-activate Out of Service Equipment Q1.P1.P3(Con).P2(Con).P2(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P2(Con) Equal to Instructions Acknowledged
Equipment Problem Log - Copy RLP | 11-Re-activate Out of Service Equipment P1.P3(Con).P2(Con).P2(Con).P3(Con).P2(Con) | Re-activate Out of Service Equipment Q1.P1.P3(Con).P2(Con).P2(Con).P3(Con).P2(Con) | Standard Acceptance Criteria Re-activate Out of Service Equipment Q1.P1.P3(Con).P2(Con).P2(Con).P3(Con).P2(Con) Equal to Instructions Acknowledged
Equipment Problem Log - Copy RLP | 11-Re-activate Out of Service Equipment P1.P3(Con).P2(Con).P2(Con).P3(Con).P3(Con).P2(Con).P2(Con) | Re-activate Out of Service Equipment Q1.P1.P3(Con).P2(Con).P2(Con).P3(Con).P3(Con).P2(Con).P2(Con) | Standard Acceptance Criteria Re-activate Out of Service Equipment Q1.P1.P3(Con).P2(Con).P2(Con).P3(Con).P3(Con).P2(Con).P2(Con) Equal to Instructions Acknowledged
Equipment Problem Log - Copy RLP | 11-Re-activate Out of Service Equipment P1.P3(Con).P2(Con).P2(Con).P3(Con).P3(Con).P2(Con).P3(Con).P2(Con).P2(Con) | Re-activate Out of Service Equipment Q1.P1.P3(Con).P2(Con).P2(Con).P3(Con).P3(Con).P2(Con).P3(Con).P2(Con).P2(Con) | Standard Acceptance Criteria Re-activate Out of Service Equipment Q1.P1.P3(Con).P2(Con).P2(Con).P3(Con).P3(Con).P2(Con).P3(Con).P2(Con).P2(Con) Equal to Instructions Acknowledged
Equipment Problem Log - Copy RLP | 11-Re-activate Out of Service Equipment P1.P3(Con).P2(Con).P2(Con).P3(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P2(Con) | Re-activate Out of Service Equipment Q1.P1.P3(Con).P2(Con).P2(Con).P3(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P2(Con) | Standard Acceptance Criteria Re-activate Out of Service Equipment Q1.P1.P3(Con).P2(Con).P2(Con).P3(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P2(Con) Equal to Instructions Acknowledged
Equipment Problem Log - Copy RLP | 11-Re-activate Out of Service Equipment P1.P3(Con).P2(Con).P2(Con).P3(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P2(Con) | Re-activate Out of Service Equipment Q1.P1.P3(Con).P2(Con).P2(Con).P3(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P2(Con) | Standard Acceptance Criteria Re-activate Out of Service Equipment Q1.P1.P3(Con).P2(Con).P2(Con).P3(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P2(Con) Equal to Instructions Acknowledged
Equipment Problem Log - Copy RLP | 11-Re-activate Out of Service Equipment P1.P3(Con).P2(Con).P2(Con).P3(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P2(Con) | Re-activate Out of Service Equipment Q1.P1.P3(Con).P2(Con).P2(Con).P3(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P2(Con) | Standard Acceptance Criteria Re-activate Out of Service Equipment Q1.P1.P3(Con).P2(Con).P2(Con).P3(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P2(Con) Equal to Instructions Acknowledged
Equipment Problem Log - Copy RLP | 11-Re-activate Out of Service Equipment P1.P3(Con).P2(Con).P2(Con).P3(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P2(Con) | Re-activate Out of Service Equipment Q1.P1.P3(Con).P2(Con).P2(Con).P3(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P2(Con) | Standard Acceptance Criteria Re-activate Out of Service Equipment Q1.P1.P3(Con).P2(Con).P2(Con).P3(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P2(Con) Equal to Instructions Acknowledged
Equipment Problem Log - Copy RLP | 11-Re-activate Out of Service Equipment P1.P3(Con).P2(Con).P2(Con).P3(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P2(Con) | Re-activate Out of Service Equipment Q1.P1.P3(Con).P2(Con).P2(Con).P3(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P2(Con) | Standard Acceptance Criteria Re-activate Out of Service Equipment Q1.P1.P3(Con).P2(Con).P2(Con).P3(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P2(Con) Equal to Instructions Acknowledged
Equipment Problem Log - Copy RLP | 11-Re-activate Out of Service Equipment P1.P3(Con).P2(Con).P2(Con).P3(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con) | Re-activate Out of Service Equipment Q1.P1.P3(Con).P2(Con).P2(Con).P3(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con) | Standard Acceptance Criteria Re-activate Out of Service Equipment Q1.P1.P3(Con).P2(Con).P2(Con).P3(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con) Equal to Instructions Acknowledged
Equipment Problem Log - Copy RLP | 11-Re-activate Out of Service Equipment P1.P3(Con).P3(Con).P2(Con).P2(Con).P2(Con) | Re-activate Out of Service Equipment Q1.P1.P3(Con).P3(Con).P2(Con).P2(Con).P2(Con) | Standard Acceptance Criteria Re-activate Out of Service Equipment Q1.P1.P3(Con).P3(Con).P2(Con).P2(Con).P2(Con) Equal to Instructions Acknowledged
Equipment Problem Log - Copy RLP | 11-Re-activate Out of Service Equipment P1.P3(Con).P3(Con).P2(Con).P2(Con).P3(Con).P2(Con).P2(Con) | Re-activate Out of Service Equipment Q1.P1.P3(Con).P3(Con).P2(Con).P2(Con).P3(Con).P2(Con).P2(Con) | Standard Acceptance Criteria Re-activate Out of Service Equipment Q1.P1.P3(Con).P3(Con).P2(Con).P2(Con).P3(Con).P2(Con).P2(Con) Equal to Instructions Acknowledged
Equipment Problem Log - Copy RLP | 11-Re-activate Out of Service Equipment P1.P3(Con).P3(Con).P2(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P2(Con) | Re-activate Out of Service Equipment Q1.P1.P3(Con).P3(Con).P2(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P2(Con) | Standard Acceptance Criteria Re-activate Out of Service Equipment Q1.P1.P3(Con).P3(Con).P2(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P2(Con) Equal to Instructions Acknowledged
Equipment Problem Log - Copy RLP | 11-Re-activate Out of Service Equipment P1.P3(Con).P3(Con).P2(Con).P3(Con).P2(Con).P2(Con) | Re-activate Out of Service Equipment Q1.P1.P3(Con).P3(Con).P2(Con).P3(Con).P2(Con).P2(Con) | Standard Acceptance Criteria Re-activate Out of Service Equipment Q1.P1.P3(Con).P3(Con).P2(Con).P3(Con).P2(Con).P2(Con) Equal to Instructions Acknowledged
Equipment Problem Log - Copy RLP | 11-Re-activate Out of Service Equipment P1.P3(Con).P3(Con).P3(Con).P2(Con) | Re-activate Out of Service Equipment Q1.P1.P3(Con).P3(Con).P3(Con).P2(Con) | Standard Acceptance Criteria Re-activate Out of Service Equipment Q1.P1.P3(Con).P3(Con).P3(Con).P2(Con) Equal to Instructions Acknowledged
Equipment Problem Log - Copy RLP | 11-Re-activate Out of Service Equipment P1.P3(Con).P3(Con).P3(Con).P3(Con).P2(Con).P2(Con) | Re-activate Out of Service Equipment Q1.P1.P3(Con).P3(Con).P3(Con).P3(Con).P2(Con).P2(Con) | Standard Acceptance Criteria Re-activate Out of Service Equipment Q1.P1.P3(Con).P3(Con).P3(Con).P3(Con).P2(Con).P2(Con) Equal to Instructions Acknowledged
Equipment Problem Log - Copy 2 | 11-Re-activate Out of Service Equipment P1.P2(Con).P2(Con).P2(Con) | Re-activate Out of Service Equipment Q1.P1.P2(Con).P2(Con).P2(Con) | Standard Acceptance Criteria Re-activate Out of Service Equipment Q1.P1.P2(Con).P2(Con).P2(Con) Equal to Instructions Acknowledged
Equipment Problem Log - Copy 2 | 11-Re-activate Out of Service Equipment P1.P2(Con).P2(Con).P3(Con).P2(Con).P2(Con) | Re-activate Out of Service Equipment Q1.P1.P2(Con).P2(Con).P3(Con).P2(Con).P2(Con) | Standard Acceptance Criteria Re-activate Out of Service Equipment Q1.P1.P2(Con).P2(Con).P3(Con).P2(Con).P2(Con) Equal to Instructions Acknowledged
Equipment Problem Log - Copy 2 | 11-Re-activate Out of Service Equipment P1.P2(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P2(Con) | Re-activate Out of Service Equipment Q1.P1.P2(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P2(Con) | Standard Acceptance Criteria Re-activate Out of Service Equipment Q1.P1.P2(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P2(Con) Equal to Instructions Acknowledged
Equipment Problem Log - Copy 2 | 11-Re-activate Out of Service Equipment P1.P2(Con).P4(Con).P2(Con).P2(Con).P2(Con).P4(Con).P2(Con) | Re-activate Out of Service Equipment Q1.P1.P2(Con).P4(Con).P2(Con).P2(Con).P2(Con).P4(Con).P2(Con) | Standard Acceptance Criteria Re-activate Out of Service Equipment Q1.P1.P2(Con).P4(Con).P2(Con).P2(Con).P2(Con).P4(Con).P2(Con) Equal to Instructions Acknowledged
Equipment Problem Log - Copy 2 | 11-Re-activate Out of Service Equipment P1.P2(Con).P4(Con).P2(Con).P2(Con).P2(Con).P4(Con).P3(Con).P2(Con).P2(Con) | Re-activate Out of Service Equipment Q1.P1.P2(Con).P4(Con).P2(Con).P2(Con).P2(Con).P4(Con).P3(Con).P2(Con).P2(Con) | Standard Acceptance Criteria Re-activate Out of Service Equipment Q1.P1.P2(Con).P4(Con).P2(Con).P2(Con).P2(Con).P4(Con).P3(Con).P2(Con).P2(Con) Equal to Instructions Acknowledged
Equipment Problem Log - Copy 2 | 11-Re-activate Out of Service Equipment P1.P2(Con).P4(Con).P2(Con).P2(Con).P2(Con).P5(Con).P2(Con).P2(Con) | Re-activate Out of Service Equipment Q1.P1.P2(Con).P4(Con).P2(Con).P2(Con).P2(Con).P5(Con).P2(Con).P2(Con) | Standard Acceptance Criteria Re-activate Out of Service Equipment Q1.P1.P2(Con).P4(Con).P2(Con).P2(Con).P2(Con).P5(Con).P2(Con).P2(Con) Equal to Instructions Acknowledged
Equipment Problem Log - Copy 2 | 11-Re-activate Out of Service Equipment P1.P2(Con).P4(Con).P2(Con).P2(Con).P2(Con).P5(Con).P2(Con).P3(Con).P2(Con).P2(Con) | Re-activate Out of Service Equipment Q1.P1.P2(Con).P4(Con).P2(Con).P2(Con).P2(Con).P5(Con).P2(Con).P3(Con).P2(Con).P2(Con) | Standard Acceptance Criteria Re-activate Out of Service Equipment Q1.P1.P2(Con).P4(Con).P2(Con).P2(Con).P2(Con).P5(Con).P2(Con).P3(Con).P2(Con).P2(Con) Equal to Instructions Acknowledged
Equipment Problem Log - Copy 2 | 11-Re-activate Out of Service Equipment P1.P2(Con).P4(Con).P2(Con).P2(Con).P2(Con).P5(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con) | Re-activate Out of Service Equipment Q1.P1.P2(Con).P4(Con).P2(Con).P2(Con).P2(Con).P5(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con) | Standard Acceptance Criteria Re-activate Out of Service Equipment Q1.P1.P2(Con).P4(Con).P2(Con).P2(Con).P2(Con).P5(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con) Equal to Instructions Acknowledged
Equipment Problem Log - Copy 2 | 11-Re-activate Out of Service Equipment P1.P2(Con).P4(Con).P2(Con).P2(Con).P3(Con).P2(Con) | Re-activate Out of Service Equipment Q1.P1.P2(Con).P4(Con).P2(Con).P2(Con).P3(Con).P2(Con) | Standard Acceptance Criteria Re-activate Out of Service Equipment Q1.P1.P2(Con).P4(Con).P2(Con).P2(Con).P3(Con).P2(Con) Equal to Instructions Acknowledged
Equipment Problem Log - Copy 2 | 11-Re-activate Out of Service Equipment P1.P2(Con).P4(Con).P2(Con).P2(Con).P3(Con).P3(Con).P2(Con).P2(Con) | Re-activate Out of Service Equipment Q1.P1.P2(Con).P4(Con).P2(Con).P2(Con).P3(Con).P3(Con).P2(Con).P2(Con) | Standard Acceptance Criteria Re-activate Out of Service Equipment Q1.P1.P2(Con).P4(Con).P2(Con).P2(Con).P3(Con).P3(Con).P2(Con).P2(Con) Equal to Instructions Acknowledged
Equipment Problem Log - Copy 2 | 11-Re-activate Out of Service Equipment P1.P2(Con).P4(Con).P3(Con).P2(Con).P2(Con).P2(Con) | Re-activate Out of Service Equipment Q1.P1.P2(Con).P4(Con).P3(Con).P2(Con).P2(Con).P2(Con) | Standard Acceptance Criteria Re-activate Out of Service Equipment Q1.P1.P2(Con).P4(Con).P3(Con).P2(Con).P2(Con).P2(Con) Equal to Instructions Acknowledged
Equipment Problem Log - Copy 2 | 11-Re-activate Out of Service Equipment P1.P2(Con).P4(Con).P3(Con).P2(Con).P2(Con).P3(Con).P2(Con).P2(Con) | Re-activate Out of Service Equipment Q1.P1.P2(Con).P4(Con).P3(Con).P2(Con).P2(Con).P3(Con).P2(Con).P2(Con) | Standard Acceptance Criteria Re-activate Out of Service Equipment Q1.P1.P2(Con).P4(Con).P3(Con).P2(Con).P2(Con).P3(Con).P2(Con).P2(Con) Equal to Instructions Acknowledged
Equipment Problem Log - Copy 2 | 11-Re-activate Out of Service Equipment P1.P2(Con).P4(Con).P3(Con).P2(Con).P3(Con).P2(Con).P2(Con) | Re-activate Out of Service Equipment Q1.P1.P2(Con).P4(Con).P3(Con).P2(Con).P3(Con).P2(Con).P2(Con) | Standard Acceptance Criteria Re-activate Out of Service Equipment Q1.P1.P2(Con).P4(Con).P3(Con).P2(Con).P3(Con).P2(Con).P2(Con) Equal to Instructions Acknowledged
Equipment Problem Log - Copy 2 | 11-Re-activate Out of Service Equipment P1.P2(Con).P4(Con).P3(Con).P3(Con).P2(Con) | Re-activate Out of Service Equipment Q1.P1.P2(Con).P4(Con).P3(Con).P3(Con).P2(Con) | Standard Acceptance Criteria Re-activate Out of Service Equipment Q1.P1.P2(Con).P4(Con).P3(Con).P3(Con).P2(Con) Equal to Instructions Acknowledged
Equipment Problem Log - Copy 2 | 11-Re-activate Out of Service Equipment P1.P2(Con).P4(Con).P3(Con).P3(Con).P3(Con).P2(Con).P2(Con) | Re-activate Out of Service Equipment Q1.P1.P2(Con).P4(Con).P3(Con).P3(Con).P3(Con).P2(Con).P2(Con) | Standard Acceptance Criteria Re-activate Out of Service Equipment Q1.P1.P2(Con).P4(Con).P3(Con).P3(Con).P3(Con).P2(Con).P2(Con) Equal to Instructions Acknowledged
Equipment Problem Log - Copy 2 | 11-Re-activate Out of Service Equipment P1.P3(Con).P2(Con).P2(Con).P2(Con).P2(Con).P2(Con) | Re-activate Out of Service Equipment Q1.P1.P3(Con).P2(Con).P2(Con).P2(Con).P2(Con).P2(Con) | Standard Acceptance Criteria Re-activate Out of Service Equipment Q1.P1.P3(Con).P2(Con).P2(Con).P2(Con).P2(Con).P2(Con) Equal to Instructions Acknowledged
Equipment Problem Log - Copy 2 | 11-Re-activate Out of Service Equipment P1.P3(Con).P2(Con).P2(Con).P2(Con).P2(Con).P3(Con).P2(Con).P2(Con) | Re-activate Out of Service Equipment Q1.P1.P3(Con).P2(Con).P2(Con).P2(Con).P2(Con).P3(Con).P2(Con).P2(Con) | Standard Acceptance Criteria Re-activate Out of Service Equipment Q1.P1.P3(Con).P2(Con).P2(Con).P2(Con).P2(Con).P3(Con).P2(Con).P2(Con) Equal to Instructions Acknowledged
Equipment Problem Log - Copy 2 | 11-Re-activate Out of Service Equipment P1.P3(Con).P2(Con).P2(Con).P2(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P2(Con) | Re-activate Out of Service Equipment Q1.P1.P3(Con).P2(Con).P2(Con).P2(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P2(Con) | Standard Acceptance Criteria Re-activate Out of Service Equipment Q1.P1.P3(Con).P2(Con).P2(Con).P2(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P2(Con) Equal to Instructions Acknowledged
Equipment Problem Log - Copy 2 | 11-Re-activate Out of Service Equipment P1.P3(Con).P2(Con).P2(Con).P2(Con).P3(Con).P2(Con).P2(Con) | Re-activate Out of Service Equipment Q1.P1.P3(Con).P2(Con).P2(Con).P2(Con).P3(Con).P2(Con).P2(Con) | Standard Acceptance Criteria Re-activate Out of Service Equipment Q1.P1.P3(Con).P2(Con).P2(Con).P2(Con).P3(Con).P2(Con).P2(Con) Equal to Instructions Acknowledged
Equipment Problem Log - Copy 2 | 11-Re-activate Out of Service Equipment P1.P3(Con).P2(Con).P2(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P2(Con) | Re-activate Out of Service Equipment Q1.P1.P3(Con).P2(Con).P2(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P2(Con) | Standard Acceptance Criteria Re-activate Out of Service Equipment Q1.P1.P3(Con).P2(Con).P2(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P2(Con) Equal to Instructions Acknowledged
Equipment Problem Log - Copy 2 | 11-Re-activate Out of Service Equipment P1.P3(Con).P2(Con).P2(Con).P3(Con).P2(Con) | Re-activate Out of Service Equipment Q1.P1.P3(Con).P2(Con).P2(Con).P3(Con).P2(Con) | Standard Acceptance Criteria Re-activate Out of Service Equipment Q1.P1.P3(Con).P2(Con).P2(Con).P3(Con).P2(Con) Equal to Instructions Acknowledged
Equipment Problem Log - Copy 2 | 11-Re-activate Out of Service Equipment P1.P3(Con).P2(Con).P2(Con).P3(Con).P3(Con).P2(Con).P2(Con) | Re-activate Out of Service Equipment Q1.P1.P3(Con).P2(Con).P2(Con).P3(Con).P3(Con).P2(Con).P2(Con) | Standard Acceptance Criteria Re-activate Out of Service Equipment Q1.P1.P3(Con).P2(Con).P2(Con).P3(Con).P3(Con).P2(Con).P2(Con) Equal to Instructions Acknowledged
Equipment Problem Log - Copy 2 | 11-Re-activate Out of Service Equipment P1.P3(Con).P2(Con).P2(Con).P3(Con).P3(Con).P2(Con).P3(Con).P2(Con).P2(Con) | Re-activate Out of Service Equipment Q1.P1.P3(Con).P2(Con).P2(Con).P3(Con).P3(Con).P2(Con).P3(Con).P2(Con).P2(Con) | Standard Acceptance Criteria Re-activate Out of Service Equipment Q1.P1.P3(Con).P2(Con).P2(Con).P3(Con).P3(Con).P2(Con).P3(Con).P2(Con).P2(Con) Equal to Instructions Acknowledged
Equipment Problem Log - Copy 2 | 11-Re-activate Out of Service Equipment P1.P3(Con).P2(Con).P2(Con).P3(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P2(Con) | Re-activate Out of Service Equipment Q1.P1.P3(Con).P2(Con).P2(Con).P3(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P2(Con) | Standard Acceptance Criteria Re-activate Out of Service Equipment Q1.P1.P3(Con).P2(Con).P2(Con).P3(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P2(Con) Equal to Instructions Acknowledged
Equipment Problem Log - Copy 2 | 11-Re-activate Out of Service Equipment P1.P3(Con).P2(Con).P2(Con).P3(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P2(Con) | Re-activate Out of Service Equipment Q1.P1.P3(Con).P2(Con).P2(Con).P3(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P2(Con) | Standard Acceptance Criteria Re-activate Out of Service Equipment Q1.P1.P3(Con).P2(Con).P2(Con).P3(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P2(Con) Equal to Instructions Acknowledged
Equipment Problem Log - Copy 2 | 11-Re-activate Out of Service Equipment P1.P3(Con).P2(Con).P2(Con).P3(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P2(Con) | Re-activate Out of Service Equipment Q1.P1.P3(Con).P2(Con).P2(Con).P3(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P2(Con) | Standard Acceptance Criteria Re-activate Out of Service Equipment Q1.P1.P3(Con).P2(Con).P2(Con).P3(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P2(Con) Equal to Instructions Acknowledged
Equipment Problem Log - Copy 2 | 11-Re-activate Out of Service Equipment P1.P3(Con).P2(Con).P2(Con).P3(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P2(Con) | Re-activate Out of Service Equipment Q1.P1.P3(Con).P2(Con).P2(Con).P3(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P2(Con) | Standard Acceptance Criteria Re-activate Out of Service Equipment Q1.P1.P3(Con).P2(Con).P2(Con).P3(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P2(Con) Equal to Instructions Acknowledged
Equipment Problem Log - Copy 2 | 11-Re-activate Out of Service Equipment P1.P3(Con).P2(Con).P2(Con).P3(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P2(Con) | Re-activate Out of Service Equipment Q1.P1.P3(Con).P2(Con).P2(Con).P3(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P2(Con) | Standard Acceptance Criteria Re-activate Out of Service Equipment Q1.P1.P3(Con).P2(Con).P2(Con).P3(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P2(Con) Equal to Instructions Acknowledged
Equipment Problem Log - Copy 2 | 11-Re-activate Out of Service Equipment P1.P3(Con).P2(Con).P2(Con).P3(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con) | Re-activate Out of Service Equipment Q1.P1.P3(Con).P2(Con).P2(Con).P3(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con) | Standard Acceptance Criteria Re-activate Out of Service Equipment Q1.P1.P3(Con).P2(Con).P2(Con).P3(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con) Equal to Instructions Acknowledged
Equipment Problem Log - Copy 2 | 11-Re-activate Out of Service Equipment P1.P3(Con).P3(Con).P2(Con).P2(Con).P2(Con) | Re-activate Out of Service Equipment Q1.P1.P3(Con).P3(Con).P2(Con).P2(Con).P2(Con) | Standard Acceptance Criteria Re-activate Out of Service Equipment Q1.P1.P3(Con).P3(Con).P2(Con).P2(Con).P2(Con) Equal to Instructions Acknowledged
Equipment Problem Log - Copy 2 | 11-Re-activate Out of Service Equipment P1.P3(Con).P3(Con).P2(Con).P2(Con).P3(Con).P2(Con).P2(Con) | Re-activate Out of Service Equipment Q1.P1.P3(Con).P3(Con).P2(Con).P2(Con).P3(Con).P2(Con).P2(Con) | Standard Acceptance Criteria Re-activate Out of Service Equipment Q1.P1.P3(Con).P3(Con).P2(Con).P2(Con).P3(Con).P2(Con).P2(Con) Equal to Instructions Acknowledged
Equipment Problem Log - Copy 2 | 11-Re-activate Out of Service Equipment P1.P3(Con).P3(Con).P2(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P2(Con) | Re-activate Out of Service Equipment Q1.P1.P3(Con).P3(Con).P2(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P2(Con) | Standard Acceptance Criteria Re-activate Out of Service Equipment Q1.P1.P3(Con).P3(Con).P2(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P2(Con) Equal to Instructions Acknowledged
Equipment Problem Log - Copy 2 | 11-Re-activate Out of Service Equipment P1.P3(Con).P3(Con).P2(Con).P3(Con).P2(Con).P2(Con) | Re-activate Out of Service Equipment Q1.P1.P3(Con).P3(Con).P2(Con).P3(Con).P2(Con).P2(Con) | Standard Acceptance Criteria Re-activate Out of Service Equipment Q1.P1.P3(Con).P3(Con).P2(Con).P3(Con).P2(Con).P2(Con) Equal to Instructions Acknowledged
Equipment Problem Log - Copy 2 | 11-Re-activate Out of Service Equipment P1.P3(Con).P3(Con).P3(Con).P2(Con) | Re-activate Out of Service Equipment Q1.P1.P3(Con).P3(Con).P3(Con).P2(Con) | Standard Acceptance Criteria Re-activate Out of Service Equipment Q1.P1.P3(Con).P3(Con).P3(Con).P2(Con) Equal to Instructions Acknowledged
Equipment Problem Log - Copy 2 | 11-Re-activate Out of Service Equipment P1.P3(Con).P3(Con).P3(Con).P3(Con).P2(Con).P2(Con) | Re-activate Out of Service Equipment Q1.P1.P3(Con).P3(Con).P3(Con).P3(Con).P2(Con).P2(Con) | Standard Acceptance Criteria Re-activate Out of Service Equipment Q1.P1.P3(Con).P3(Con).P3(Con).P3(Con).P2(Con).P2(Con) Equal to Instructions Acknowledged
Bioarray Daily As Needed | QIAGEN QIAcube Daily Maintenance P1 | Remove Labware, Samples and Reagents Q1.P1 | Standard Acceptance Criteria Remove Labware, Samples and Reagents Q1.P1 Equal to Performed - OK
Bioarray Daily As Needed | QIAGEN QIAcube Daily Maintenance P1 | Recap Reagent Bottles Q2.P1 | Standard Acceptance Criteria Recap Reagent Bottles Q2.P1 Equal to Performed - OK
Bioarray Daily As Needed | QIAGEN QIAcube Daily Maintenance P1 | Empty Waste and Clean Q3.P1 | Standard Acceptance Criteria Empty Waste and Clean Q3.P1 Equal to Performed - OK
Bioarray Daily As Needed | QIAGEN QIAcube Daily Maintenance P1 | Clean Shaker Rack and Bottle Rack Q4.P1 | Standard Acceptance Criteria Clean Shaker Rack and Bottle Rack Q4.P1 Equal to Performed - OK
QIAcube Weekly | QIAGEN QIAcube Weekly Maintenance P1 | Perform QIAGEN Daily Maintenance Q1.P1 | Standard Acceptance Criteria Perform QIAGEN Daily Maintenance Q1.P1 Equal to Performed - OK
QIAcube Weekly | QIAGEN QIAcube Weekly Maintenance P1 | Clean Optical Sensor Q2.P1 | Standard Acceptance Criteria Clean Optical Sensor Q2.P1 Equal to Performed - OK
QIAcube Weekly | QIAGEN QIAcube Weekly Maintenance P1 | Clean Liner of Waste Drawer Q3.P1 | Standard Acceptance Criteria Clean Liner of Waste Drawer Q3.P1 Equal to Performed - OK
QIAcube Weekly | QIAGEN QIAcube Weekly Maintenance P1 | Clean Worktable Q4.P1 | Standard Acceptance Criteria Clean Worktable Q4.P1 Equal to Performed - OK
Bioarray Maintenance Bi-Annual | QIAGEN QIAcube Biannual Maintenance P1 | Clean QIAGEN Parts and Accessories Q1.P1 | Standard Acceptance Criteria Clean QIAGEN Parts and Accessories Q1.P1 Equal to Performed - OK
Bioarray Maintenance Bi-Annual | QIAGEN QIAcube Biannual Maintenance P1 | Clean Centrifuge Interior Q2.P1 | Standard Acceptance Criteria Clean Centrifuge Interior Q2.P1 Equal to Performed - OK
Bioarray Maintenance Bi-Annual | QIAGEN QIAcube Biannual Maintenance P1 | Centrifuge Gasket Damaged? Q3.P1 | Standard Acceptance Criteria Centrifuge Gasket Damaged? Q3.P1 Equal to No
Bioarray Maintenance Bi-Annual | QIAGEN QIAcube Biannual Maintenance P1 | Perform Tightness Test Q4.P1 | Standard Acceptance Criteria Perform Tightness Test Q4.P1 Equal to Performed - OK
BioArray Extraction Area Wipe Test Monthly | BioArray Extraction Area Wipe Test Monthly P1 | Final Result of Wipe Test for Extraction Area? Q1.P1 | Standard Acceptance Criteria Final Result of Wipe Test for Extraction Area? Q1.P1 Equal to NTD (No Type Determined) - Acceptable
Hemocue Shutdown Daily | Shutdown P1 | Clean and Shut Down Instrument? Q1.P1 | Standard Acceptance Criteria Clean and Shut Down Instrument? Q1.P1 Equal to Performed - OK
Hemocue Shutdown Daily | Shutdown P1 | Problem with Equipment Q2.P1 | Standard Acceptance Criteria Problem with Equipment Q2.P1 Equal to No
QIAcube Monthly | QIAcube Monthly QC P1 | Clean Optical Sensor Q1.P1 | Standard Acceptance Criteria Clean Optical Sensor Q1.P1 Equal to Performed - OK
QIAcube Monthly | QIAcube Monthly QC P1 | Clean Liner of Waste Drawer Q2.P1 | Standard Acceptance Criteria Clean Liner of Waste Drawer Q2.P1 Equal to Performed - OK
QIAcube Monthly | QIAcube Monthly QC P1 | Clean Worktable Q3.P1 | Standard Acceptance Criteria Clean Worktable Q3.P1 Equal to Performed - OK
Bioarray AIS400 Weekly | AIS400 Weekly Maintenance P1 | Clean AIS400 Q1.P1 | Standard Acceptance Criteria Clean AIS400 Q1.P1 Equal to Performed - OK
Bioarray Microcentrifuge Monthly | BioArray Denville Monthly P1 | Clean Denville Microcentrifuge Q1.P1 | Standard Acceptance Criteria Clean Denville Microcentrifuge Q1.P1 Equal to Performed - OK
Sorvall Cell Washer Calibration Annual | Annual Sorvall CW Calibration Final Results P18 | Calibration Acceptable Q4.P18 | Standard Acceptance Criteria Calibration Acceptable Q4.P18 Equal to Acceptable
Neo Correlation Reflex ABO Testing | ABORH and ReflexABO (Automated) P1 | Results Acceptable? Q22.P1 | Standard Acceptance Criteria Results Acceptable? Q22.P1 Equal to Acceptable
Neo Correlation Reflex ABO Testing | ABORH and ReflexABO (Automated) P1 | Problem with Equipment Q23.P1 | Standard Acceptance Criteria Problem with Equipment Q23.P1 Equal to No
Neo Correlation Antibody Testing | 2 Cell Antibody Testing P1.P2(Con) | Results Acceptable? Q16.P1.P2(Con) | Standard Acceptance Criteria Results Acceptable? Q16.P1.P2(Con) Equal to Acceptable
Neo Correlation Antibody Testing | 2 Cell Antibody Testing P1.P2(Con) | Problem with Equipment Q17.P1.P2(Con) | Standard Acceptance Criteria Problem with Equipment Q17.P1.P2(Con) Equal to No
Neo Correlation Antibody Testing | Pool Cell Antibody Testing P1.P3(Con) | Results Acceptable? Q16.P1.P3(Con) | Standard Acceptance Criteria Results Acceptable? Q16.P1.P3(Con) Equal to Acceptable
Neo Correlation Antibody Testing | Pool Cell Antibody Testing P1.P3(Con) | Problem with Equipment Q17.P1.P3(Con) | Standard Acceptance Criteria Problem with Equipment Q17.P1.P3(Con) Equal to No
Neo Correlation Weak D Testing | Weak D P1 | Results Acceptable? Q16.P1 | Standard Acceptance Criteria Results Acceptable? Q16.P1 Equal to Acceptable
Neo Correlation Weak D Testing | Weak D P1 | Problem with Equipment Q17.P1 | Standard Acceptance Criteria Problem with Equipment Q17.P1 Equal to No
Neo Correlation CMV Testing | CMV P1 | Results Acceptable? Q5.P1 | Standard Acceptance Criteria Results Acceptable? Q5.P1 Equal to Acceptable
Neo Correlation CMV Testing | CMV P1 | Results Acceptable? Q10.P1 | Standard Acceptance Criteria Results Acceptable? Q10.P1 Equal to Acceptable
Neo Correlation CMV Testing | CMV P1 | Results Acceptable? Q15.P1 | Standard Acceptance Criteria Results Acceptable? Q15.P1 Equal to Acceptable
Neo Correlation CMV Testing | CMV P1 | Problem with Equipment Q16.P1 | Standard Acceptance Criteria Problem with Equipment Q16.P1 Equal to No
Hematology Analyzer Parallels Bi-Annual | Hematology Bi-Annual Parallels Instrument 1 P1 | Sample 1 WBC Value Q2.P1 | Standard Acceptance Criteria Sample 1 WBC Value Q2.P1 Between 0 to 999
Hematology Analyzer Parallels Bi-Annual | Hematology Bi-Annual Parallels Instrument 1 P1 | Sample 1 RBC Value Q3.P1 | Standard Acceptance Criteria Sample 1 RBC Value Q3.P1 Between 0 to 999
Hematology Analyzer Parallels Bi-Annual | Hematology Bi-Annual Parallels Instrument 1 P1 | Sample 1 HGB Value Q4.P1 | Standard Acceptance Criteria Sample 1 HGB Value Q4.P1 Between 0 to 999
Hematology Analyzer Parallels Bi-Annual | Hematology Bi-Annual Parallels Instrument 1 P1 | Sample 1 HCT Value Q5.P1 | Standard Acceptance Criteria Sample 1 HCT Value Q5.P1 Between 0 to 999
Hematology Analyzer Parallels Bi-Annual | Hematology Bi-Annual Parallels Instrument 1 P1 | Sample 1 PLT Value Q6.P1 | Standard Acceptance Criteria Sample 1 PLT Value Q6.P1 Between 0 to 999
Hematology Analyzer Parallels Bi-Annual | Hematology Bi-Annual Parallels Instrument 2 P2 | Sample 1 WBC Value Q2.P2 | Standard Acceptance Criteria Sample 1 WBC Value Q2.P2 Between 0 to 999
Hematology Analyzer Parallels Bi-Annual | Hematology Bi-Annual Parallels Instrument 2 P2 | Sample 1 RBC Value Q3.P2 | Standard Acceptance Criteria Sample 1 RBC Value Q3.P2 Between 0 to 999
Hematology Analyzer Parallels Bi-Annual | Hematology Bi-Annual Parallels Instrument 2 P2 | Sample 1 HGB Value Q4.P2 | Standard Acceptance Criteria Sample 1 HGB Value Q4.P2 Between 0 to 999
Hematology Analyzer Parallels Bi-Annual | Hematology Bi-Annual Parallels Instrument 2 P2 | Sample 1 HCT Value Q5.P2 | Standard Acceptance Criteria Sample 1 HCT Value Q5.P2 Between 0 to 999
Hematology Analyzer Parallels Bi-Annual | Hematology Bi-Annual Parallels Instrument 2 P2 | Sample 1 PLT Value Q6.P2 | Standard Acceptance Criteria Sample 1 PLT Value Q6.P2 Between 0 to 999
Hematology Analyzer Parallels Bi-Annual | Hematology Bi-Annual Parallels Instrument 3 P3 | Sample 1 WBC Value Q2.P3 | Standard Acceptance Criteria Sample 1 WBC Value Q2.P3 Between 0 to 999
Hematology Analyzer Parallels Bi-Annual | Hematology Bi-Annual Parallels Instrument 3 P3 | Sample 1 RBC Value Q3.P3 | Standard Acceptance Criteria Sample 1 RBC Value Q3.P3 Between 0 to 999
Hematology Analyzer Parallels Bi-Annual | Hematology Bi-Annual Parallels Instrument 3 P3 | Sample 1 HGB Value Q4.P3 | Standard Acceptance Criteria Sample 1 HGB Value Q4.P3 Between 0 to 999
Hematology Analyzer Parallels Bi-Annual | Hematology Bi-Annual Parallels Instrument 3 P3 | Sample 1 HCT Value Q5.P3 | Standard Acceptance Criteria Sample 1 HCT Value Q5.P3 Between 0 to 999
Hematology Analyzer Parallels Bi-Annual | Hematology Bi-Annual Parallels Instrument 3 P3 | Sample 1 PLT Value Q6.P3 | Standard Acceptance Criteria Sample 1 PLT Value Q6.P3 Between 0 to 999
Hematology Analyzer Parallels Bi-Annual | Hematology Bi-Annual Parallels Instrument 4 P4 | Sample 1 WBC Value Q2.P4 | Standard Acceptance Criteria Sample 1 WBC Value Q2.P4 Between 0 to 999
Hematology Analyzer Parallels Bi-Annual | Hematology Bi-Annual Parallels Instrument 4 P4 | Sample 1 RBC Value Q3.P4 | Standard Acceptance Criteria Sample 1 RBC Value Q3.P4 Between 0 to 999
Hematology Analyzer Parallels Bi-Annual | Hematology Bi-Annual Parallels Instrument 4 P4 | Sample 1 HGB Value Q4.P4 | Standard Acceptance Criteria Sample 1 HGB Value Q4.P4 Between 0 to 999
Hematology Analyzer Parallels Bi-Annual | Hematology Bi-Annual Parallels Instrument 4 P4 | Sample 1 HCT Value Q5.P4 | Standard Acceptance Criteria Sample 1 HCT Value Q5.P4 Between 0 to 999
Hematology Analyzer Parallels Bi-Annual | Hematology Bi-Annual Parallels Instrument 4 P4 | Sample 1 PLT Value Q6.P4 | Standard Acceptance Criteria Sample 1 PLT Value Q6.P4 Between 0 to 999
Hematology Analyzer Parallels Bi-Annual | Hematology Bi-Annual Parallels Instrument 5 P5 | Sample 1 WBC Value Q2.P5 | Standard Acceptance Criteria Sample 1 WBC Value Q2.P5 Between 0 to 999
Hematology Analyzer Parallels Bi-Annual | Hematology Bi-Annual Parallels Instrument 5 P5 | Sample 1 RBC Value Q3.P5 | Standard Acceptance Criteria Sample 1 RBC Value Q3.P5 Between 0 to 999
Hematology Analyzer Parallels Bi-Annual | Hematology Bi-Annual Parallels Instrument 5 P5 | Sample 1 HGB Value Q4.P5 | Standard Acceptance Criteria Sample 1 HGB Value Q4.P5 Between 0 to 999
Hematology Analyzer Parallels Bi-Annual | Hematology Bi-Annual Parallels Instrument 5 P5 | Sample 1 HCT Value Q5.P5 | Standard Acceptance Criteria Sample 1 HCT Value Q5.P5 Between 0 to 999
Hematology Analyzer Parallels Bi-Annual | Hematology Bi-Annual Parallels Instrument 5 P5 | Sample 1 PLT Value Q6.P5 | Standard Acceptance Criteria Sample 1 PLT Value Q6.P5 Between 0 to 999
Hematology Analyzer Parallels Bi-Annual | Hematology Bi-Annual Parallels Acceptance P6 | Are Sample 1's WBC Results from All 4 Equipment's Within +/= 10% of Each Other? Q1.P6 | Standard Acceptance Criteria Are Sample 1's WBC Results from All 4 Equipment's Within +/= 10% of Each Other? Q1.P6 Equal to Yes
Hematology Analyzer Parallels Bi-Annual | Hematology Bi-Annual Parallels Acceptance P6 | Are Sample 1's RBC Results from All 4 Equipment's Within +/= 10% of Each Other? Q2.P6 | Standard Acceptance Criteria Are Sample 1's RBC Results from All 4 Equipment's Within +/= 10% of Each Other? Q2.P6 Equal to Yes
Hematology Analyzer Parallels Bi-Annual | Hematology Bi-Annual Parallels Acceptance P6 | Are Sample 1's HGB Results from All 4 Equipment's Within +/= 10% of Each Other? Q3.P6 | Standard Acceptance Criteria Are Sample 1's HGB Results from All 4 Equipment's Within +/= 10% of Each Other? Q3.P6 Equal to Yes
Hematology Analyzer Parallels Bi-Annual | Hematology Bi-Annual Parallels Acceptance P6 | Are Sample 1's HCT Results from All 4 Equipment's Within +/= 10% of Each Other? Q4.P6 | Standard Acceptance Criteria Are Sample 1's HCT Results from All 4 Equipment's Within +/= 10% of Each Other? Q4.P6 Equal to Yes
Hematology Analyzer Parallels Bi-Annual | Hematology Bi-Annual Parallels Acceptance P6 | Are Sample 1's PLT Results from All 4 Equipment's Within +/= 10% of Each Other? Q5.P6 | Standard Acceptance Criteria Are Sample 1's PLT Results from All 4 Equipment's Within +/= 10% of Each Other? Q5.P6 Equal to Yes
Hematology Analyzer Parallels Bi-Annual | Hematology Bi-Annual Parallels Acceptance P6 | Are Sample 2's WBC Results from All 4 Equipment's Within +/= 10% of Each Other? Q6.P6 | Standard Acceptance Criteria Are Sample 2's WBC Results from All 4 Equipment's Within +/= 10% of Each Other? Q6.P6 Equal to Yes
Hematology Analyzer Parallels Bi-Annual | Hematology Bi-Annual Parallels Acceptance P6 | Are Sample 2's RBC Results from All 4 Equipment's Within +/= 10% of Each Other? Q7.P6 | Standard Acceptance Criteria Are Sample 2's RBC Results from All 4 Equipment's Within +/= 10% of Each Other? Q7.P6 Equal to Yes
Hematology Analyzer Parallels Bi-Annual | Hematology Bi-Annual Parallels Acceptance P6 | Are Sample 2's HGB Results from All 4 Equipment's Within +/= 10% of Each Other? Q8.P6 | Standard Acceptance Criteria Are Sample 2's HGB Results from All 4 Equipment's Within +/= 10% of Each Other? Q8.P6 Equal to Yes
Hematology Analyzer Parallels Bi-Annual | Hematology Bi-Annual Parallels Acceptance P6 | Are Sample 2's HCT Results from All 4 Equipment's Within +/= 10% of Each Other? Q9.P6 | Standard Acceptance Criteria Are Sample 2's HCT Results from All 4 Equipment's Within +/= 10% of Each Other? Q9.P6 Equal to Yes
Hematology Analyzer Parallels Bi-Annual | Hematology Bi-Annual Parallels Acceptance P6 | Are Sample 2's PLT Results from All 4 Equipment's Within +/= 10% of Each Other? Q10.P6 | Standard Acceptance Criteria Are Sample 2's PLT Results from All 4 Equipment's Within +/= 10% of Each Other? Q10.P6 Equal to Yes
Hematology Analyzer Parallels Bi-Annual | Hematology Bi-Annual Parallels Acceptance P6 | Are Sample 3's WBC Results from All 4 Equipment's Within +/= 10% of Each Other? Q11.P6 | Standard Acceptance Criteria Are Sample 3's WBC Results from All 4 Equipment's Within +/= 10% of Each Other? Q11.P6 Equal to Yes
Hematology Analyzer Parallels Bi-Annual | Hematology Bi-Annual Parallels Acceptance P6 | Are Sample 3's RBC Results from All 4 Equipment's Within +/= 10% of Each Other? Q12.P6 | Standard Acceptance Criteria Are Sample 3's RBC Results from All 4 Equipment's Within +/= 10% of Each Other? Q12.P6 Equal to Yes
Hematology Analyzer Parallels Bi-Annual | Hematology Bi-Annual Parallels Acceptance P6 | Are Sample 3's HGB Results from All 4 Equipment's Within +/= 10% of Each Other? Q13.P6 | Standard Acceptance Criteria Are Sample 3's HGB Results from All 4 Equipment's Within +/= 10% of Each Other? Q13.P6 Equal to Yes
Hematology Analyzer Parallels Bi-Annual | Hematology Bi-Annual Parallels Acceptance P6 | Are Sample 3's HCT Results from All 4 Equipment's Within +/= 10% of Each Other? Q14.P6 | Standard Acceptance Criteria Are Sample 3's HCT Results from All 4 Equipment's Within +/= 10% of Each Other? Q14.P6 Equal to Yes
Hematology Analyzer Parallels Bi-Annual | Hematology Bi-Annual Parallels Acceptance P6 | Are Sample 3's PLT Results from All 4 Equipment's Within +/= 10% of Each Other? Q15.P6 | Standard Acceptance Criteria Are Sample 3's PLT Results from All 4 Equipment's Within +/= 10% of Each Other? Q15.P6 Equal to Yes
Hematology Analyzer Parallels Bi-Annual | Hematology Bi-Annual Parallels Acceptance P6 | Are Sample 4's WBC Results from All 4 Equipment's Within +/= 10% of Each Other? Q16.P6 | Standard Acceptance Criteria Are Sample 4's WBC Results from All 4 Equipment's Within +/= 10% of Each Other? Q16.P6 Equal to Yes
Hematology Analyzer Parallels Bi-Annual | Hematology Bi-Annual Parallels Acceptance P6 | Are Sample 4's RBC Results from All 4 Equipment's Within +/= 10% of Each Other? Q17.P6 | Standard Acceptance Criteria Are Sample 4's RBC Results from All 4 Equipment's Within +/= 10% of Each Other? Q17.P6 Equal to Yes
Hematology Analyzer Parallels Bi-Annual | Hematology Bi-Annual Parallels Acceptance P6 | Are Sample 4's HGB Results from All 4 Equipment's Within +/= 10% of Each Other? Q18.P6 | Standard Acceptance Criteria Are Sample 4's HGB Results from All 4 Equipment's Within +/= 10% of Each Other? Q18.P6 Equal to Yes
Hematology Analyzer Parallels Bi-Annual | Hematology Bi-Annual Parallels Acceptance P6 | Are Sample 4's HCT Results from All 4 Equipment's Within +/= 10% of Each Other? Q19.P6 | Standard Acceptance Criteria Are Sample 4's HCT Results from All 4 Equipment's Within +/= 10% of Each Other? Q19.P6 Equal to Yes
Hematology Analyzer Parallels Bi-Annual | Hematology Bi-Annual Parallels Acceptance P6 | Are Sample 4's PLT Results from All 4 Equipment's Within +/= 10% of Each Other? Q20.P6 | Standard Acceptance Criteria Are Sample 4's PLT Results from All 4 Equipment's Within +/= 10% of Each Other? Q20.P6 Equal to Yes
Hematology Analyzer Parallels Bi-Annual | Hematology Bi-Annual Parallels Acceptance P6 | Are Sample 5's WBC Results from All 4 Equipment's Within +/= 10% of Each Other? Q21.P6 | Standard Acceptance Criteria Are Sample 5's WBC Results from All 4 Equipment's Within +/= 10% of Each Other? Q21.P6 Equal to Yes
Hematology Analyzer Parallels Bi-Annual | Hematology Bi-Annual Parallels Acceptance P6 | Are Sample 5's RBC Results from All 4 Equipment's Within +/= 10% of Each Other? Q22.P6 | Standard Acceptance Criteria Are Sample 5's RBC Results from All 4 Equipment's Within +/= 10% of Each Other? Q22.P6 Equal to Yes
Hematology Analyzer Parallels Bi-Annual | Hematology Bi-Annual Parallels Acceptance P6 | Are Sample 5's HGB Results from All 4 Equipment's Within +/= 10% of Each Other? Q23.P6 | Standard Acceptance Criteria Are Sample 5's HGB Results from All 4 Equipment's Within +/= 10% of Each Other? Q23.P6 Equal to Yes
Hematology Analyzer Parallels Bi-Annual | Hematology Bi-Annual Parallels Acceptance P6 | Are Sample 5's HCT Results from All 4 Equipment's Within +/= 10% of Each Other? Q24.P6 | Standard Acceptance Criteria Are Sample 5's HCT Results from All 4 Equipment's Within +/= 10% of Each Other? Q24.P6 Equal to Yes
Hematology Analyzer Parallels Bi-Annual | Hematology Bi-Annual Parallels Acceptance P6 | Are Sample 5's PLT Results from All 4 Equipment's Within +/= 10% of Each Other? Q25.P6 | Standard Acceptance Criteria Are Sample 5's PLT Results from All 4 Equipment's Within +/= 10% of Each Other? Q25.P6 Equal to Yes
Hematology Analyzer Parallels Bi-Annual | Hematology Bi-Annual Parallels Acceptance P6 | Problem with Equipment Q27.P6 | Standard Acceptance Criteria Problem with Equipment Q27.P6 Equal to No
Hemocue Lot to Lot As Needed | Hemocue Lot to Lot P1 | Controls at Room Temperature Q12.P1 | Standard Acceptance Criteria Controls at Room Temperature Q12.P1 Equal to Yes
Hemocue Lot to Lot As Needed | Hemocue Lot to Lot P1 | Current Normal Control Open Bottle Expiration Date Q1.P1 | Standard Acceptance Criteria Current Normal Control Open Bottle Expiration Date Q1.P1 Greater than or equal to Today
Hemocue Lot to Lot As Needed | Hemocue Lot to Lot P1 | Current Abnormal High Control Open Bottle Expiration Date Q2.P1 | Standard Acceptance Criteria Current Abnormal High Control Open Bottle Expiration Date Q2.P1 Greater than or equal to Today
Hemocue Lot to Lot As Needed | Hemocue Lot to Lot P1 | Accept Current Lot Levy-Jennings Chart? Q5.P1 | Standard Acceptance Criteria Accept Current Lot Levy-Jennings Chart? Q5.P1 Equal to Acceptable
Hemocue Lot to Lot As Needed | Hemocue Lot to Lot P1 | New Normal Control Open Bottle Expiration Date Q6.P1 | Standard Acceptance Criteria New Normal Control Open Bottle Expiration Date Q6.P1 Greater than or equal to Today
Hemocue Lot to Lot As Needed | Hemocue Lot to Lot P1 | New Abnormal High Control Open Bottle Expiration Date Q7.P1 | Standard Acceptance Criteria New Abnormal High Control Open Bottle Expiration Date Q7.P1 Greater than or equal to Today
Hemocue Lot to Lot As Needed | Hemocue Lot to Lot P1 | Accept Current Lot Levy-Jennings Chart? Q10.P1 | Standard Acceptance Criteria Accept Current Lot Levy-Jennings Chart? Q10.P1 Equal to Acceptable
Hemocue Lot to Lot As Needed | Hemocue Lot to Lot P1 | Problem with QC Q11.P1 | Standard Acceptance Criteria Problem with QC Q11.P1 Equal to No
Blood Component QC Review Monthly | Monthly Blood Component QC Platelets P1 | Do 90% of the Units Tested Have a PH ≥ 6.2? Q1.P1 | Standard Acceptance Criteria Do 90% of the Units Tested Have a PH ≥ 6.2? Q1.P1 Equal to Yes
Blood Component QC Review Monthly | Monthly Blood Component QC Platelets P1 | Does 90% of the Units Tested Contain ≥ 5.5 X 1010 PLTS/unit? Q3.P1 | Standard Acceptance Criteria Does 90% of the Units Tested Contain ≥ 5.5 X 1010 PLTS/unit? Q3.P1 Equal to Yes
Blood Component QC Review Monthly | Monthly Blood Component QC Platelets P1 | Do 90% of the Units Tested Have a PH ≥ 6.2? Q5.P1 | Standard Acceptance Criteria Do 90% of the Units Tested Have a PH ≥ 6.2? Q5.P1 Equal to Yes
Blood Component QC Review Monthly | Monthly Blood Component QC Platelets P1 | Does 90% of the Units Tested Contain ≥ 3.0 X 1011 Plts/unit? (Amicus) Q7.P1 | Standard Acceptance Criteria Does 90% of the Units Tested Contain ≥ 3.0 X 1011 Plts/unit? (Amicus) Q7.P1 Equal to Yes
Blood Component QC Review Monthly | Monthly Blood Component QC Platelets P1 | Does 90% of the Units Tested Contain ≥ 3.0 X 1011 Plts/unit? (Trima) Q9.P1 | Standard Acceptance Criteria Does 90% of the Units Tested Contain ≥ 3.0 X 1011 Plts/unit? (Trima) Q9.P1 Equal to Yes
Blood Component QC Review Monthly | Monthly Blood Component QC RBCs P2 | Does 95% of the Units Tested Have a Hemoglobin >50g/unit? Q1.P2 | Standard Acceptance Criteria Does 95% of the Units Tested Have a Hemoglobin >50g/unit? Q1.P2 Equal to Yes
Blood Component QC Review Monthly | Monthly Blood Component QC Granulocytes P3 | Does 75% of the Units Tested Contain ≥ 1.0 X 1010 Granulocytes/unit? (Optia) Q1.P3 | Standard Acceptance Criteria Does 75% of the Units Tested Contain ≥ 1.0 X 1010 Granulocytes/unit? (Optia) Q1.P3 Equal to Yes
Stopwatch Calibration Annual | Stopwatch Annual Calibration P1 | Calibration Acceptable Q2.P1 | Standard Acceptance Criteria Calibration Acceptable Q2.P1 Equal to Acceptable
Stopwatch Calibration Annual | Stopwatch Annual Calibration P1 | Problem with Equipment Q4.P1 | Standard Acceptance Criteria Problem with Equipment Q4.P1 Equal to No
MT PLT XM Evening Shift Daily | Capture-P Washstation Daily P1 | Empty Waste and Clean Q1.P1 | Standard Acceptance Criteria Empty Waste and Clean Q1.P1 Equal to Performed - OK
MT PLT XM Evening Shift Daily | Capture-P Washstation Daily P1 | Check/fill PBS & DI Water Q2.P1 | Standard Acceptance Criteria Check/fill PBS & DI Water Q2.P1 Equal to Performed - OK
MT PLT XM Evening Shift Daily | Capture-P Washstation Daily P1 | Prime / Check Probe Q3.P1 | Standard Acceptance Criteria Prime / Check Probe Q3.P1 Equal to Performed - OK
MT PLT XM Evening Shift Daily | Capture-P Washstation Daily P1 | Residual Saline Verification Q4.P1 | Standard Acceptance Criteria Residual Saline Verification Q4.P1 Equal to Performed - OK
MT PLT XM Evening Shift Daily | Capture-P Washstation Daily P1 | Record Heatblock Incubator Temperature Q5.P1 | Standard Acceptance Criteria Record Heatblock Incubator Temperature Q5.P1 Between 36 to 38
MT PLT XM Evening Shift Daily | Capture-P Washstation Daily PBS P1.P2(Con) | Buffered Saline PH Result Q1.P1.P2(Con) | Standard Acceptance Criteria Buffered Saline PH Result Q1.P1.P2(Con) Between 6.8 to 7.4
MT PLT XM Evening Shift Daily | Capture-P Washstation Daily PBS P1.P2(Con) | New PBS Expiration Date Q2.P1.P2(Con) | Standard Acceptance Criteria New PBS Expiration Date Q2.P1.P2(Con) Greater than or equal to Today
Reagent Preparation Incubator Daily | QC Performed P2 | Was the QC Performed Q1.P2 | Standard Acceptance Criteria Was the QC Performed Q1.P2 Equal to Yes
Reagent Preparation Incubator Daily | Reagent Preparation Incubator P2.P2(Con) | Record Average Temperature Q1.P2.P2(Con) | Standard Acceptance Criteria Record Average Temperature Q1.P2.P2(Con) Between 30 to 34
Reagent Preparation Incubator Daily | Reagent Preparation Incubator P2.P2(Con) | Record Max Temperature Q2.P2.P2(Con) | Standard Acceptance Criteria Record Max Temperature Q2.P2.P2(Con) Less than or equal to 35
Reagent Preparation Incubator Daily | Reagent Preparation Incubator P2.P2(Con) | Problem with Equipment Q3.P2.P2(Con) | Standard Acceptance Criteria Problem with Equipment Q3.P2.P2(Con) Equal to No
RBC/Apheresis QC Data | RBC/Apheresis QC Data Results P2 | Total Hgb of Unit >50 G/dL Q1.P2 | Standard Acceptance Criteria Total Hgb of Unit >50 G/dL Q1.P2 Greater than or equal to 50
RBC/Apheresis QC Data | RBC/Apheresis QC Data Results P2 | Average Hgb ≥ 60 G/dL? Q2.P2 | Standard Acceptance Criteria Average Hgb ≥ 60 G/dL? Q2.P2 Greater than or equal to 60
RBC/Apheresis QC Data | RBC/Apheresis QC Data Results P2 | Does 95% of the Units Tested Have a Hemoglobin >50 G/dL Unit? Q3.P2 | Standard Acceptance Criteria Does 95% of the Units Tested Have a Hemoglobin >50 G/dL Unit? Q3.P2 Greater than or equal to 95
RBC/Apheresis QC Data | RBC/Apheresis QC Data Results P2 | RBC/Apheresis QC Passed? Q4.P2 | Standard Acceptance Criteria RBC/Apheresis QC Passed? Q4.P2 Equal to Yes
REES Validation/PM Annual | Annual REES Validation/PM P1 | Validation Complete Q2.P1 | Standard Acceptance Criteria Validation Complete Q2.P1 Equal to Performed
REES Validation/PM Annual | Annual REES Validation/PM P1 | Validation Reviewed Q3.P1 | Standard Acceptance Criteria Validation Reviewed Q3.P1 Equal to Performed
REES Validation/PM Annual | Annual REES Validation/PM P1 | Problem with Equipment Q5.P1 | Standard Acceptance Criteria Problem with Equipment Q5.P1 Equal to No
TS Electronic Balance Weekly | Electronic Balance Weekly P1 | Cleaned Q1.P1 | Standard Acceptance Criteria Cleaned Q1.P1 Equal to Performed - OK
TS Electronic Balance Weekly | Electronic Balance Weekly P1 | Problems with Equipment/QC Q2.P1 | Standard Acceptance Criteria Problems with Equipment/QC Q2.P1 Equal to No
Infectious Disease WNV Reagent Qualfication | Infectious Disease WNV Reagent Qualification P1 | WNV Negative External Control Result Q2.P1 | Standard Acceptance Criteria WNV Negative External Control Result Q2.P1 Equal to Negative
Infectious Disease WNV Reagent Qualfication | Infectious Disease WNV Reagent Qualification P1 | WNV Positive External Control Result Q3.P1 | Standard Acceptance Criteria WNV Positive External Control Result Q3.P1 Equal to Positive
Infectious Disease WNV Reagent Qualfication | Infectious Disease WNV Reagent Qualification P1 | NAT Calibrators and External Controls Acceptable? Q5.P1 | Standard Acceptance Criteria NAT Calibrators and External Controls Acceptable? Q5.P1 Equal to Acceptable
Infectious Disease WNV Reagent Qualfication | Infectious Disease WNV Reagent Qualification P1 | Problem with Equipment Q4.P1 | Standard Acceptance Criteria Problem with Equipment Q4.P1 Equal to No
Infectious Disease ULTRIO Reagent Qualfication | Infectious Disease ULTRIO Reagent Qualfication P1 | Ultrio Negative External Control Result Q2.P1 | Standard Acceptance Criteria Ultrio Negative External Control Result Q2.P1 Equal to Negative
Infectious Disease ULTRIO Reagent Qualfication | Infectious Disease ULTRIO Reagent Qualfication P1 | HIV-1 Positive External Control Result Q3.P1 | Standard Acceptance Criteria HIV-1 Positive External Control Result Q3.P1 Equal to Positive
Infectious Disease ULTRIO Reagent Qualfication | Infectious Disease ULTRIO Reagent Qualfication P1 | HCV Positive External Control Result Q4.P1 | Standard Acceptance Criteria HCV Positive External Control Result Q4.P1 Equal to Positive
Infectious Disease ULTRIO Reagent Qualfication | Infectious Disease ULTRIO Reagent Qualfication P1 | HBV Positive External Control Result Q5.P1 | Standard Acceptance Criteria HBV Positive External Control Result Q5.P1 Equal to Positive
Infectious Disease ULTRIO Reagent Qualfication | Infectious Disease ULTRIO Reagent Qualfication P1 | NAT Calibrators and External Controls Acceptable? Q6.P1 | Standard Acceptance Criteria NAT Calibrators and External Controls Acceptable? Q6.P1 Equal to Acceptable
Infectious Disease ULTRIO Reagent Qualfication | Infectious Disease ULTRIO Reagent Qualfication P1 | Problem with Equipment Q7.P1 | Standard Acceptance Criteria Problem with Equipment Q7.P1 Equal to No
Infectious Disease PRISM Reagent Qualfication | Infectious Disease Prism Reagent Qualfication P1 | Prism Negative External Control Channel A Result Q2.P1 | Standard Acceptance Criteria Prism Negative External Control Channel A Result Q2.P1 Equal to Negative
Infectious Disease PRISM Reagent Qualfication | Infectious Disease Prism Reagent Qualfication P1 | Prism Negative External Control Channel B Result Q3.P1 | Standard Acceptance Criteria Prism Negative External Control Channel B Result Q3.P1 Equal to Negative
Infectious Disease PRISM Reagent Qualfication | Infectious Disease Prism Reagent Qualfication P1 | Prism Positive External Control Channel A Result Q4.P1 | Standard Acceptance Criteria Prism Positive External Control Channel A Result Q4.P1 Equal to Positive
Infectious Disease PRISM Reagent Qualfication | Infectious Disease Prism Reagent Qualfication P1 | Prism Positive External Control Channel B Result Q5.P1 | Standard Acceptance Criteria Prism Positive External Control Channel B Result Q5.P1 Equal to Positive
Infectious Disease PRISM Reagent Qualfication | Infectious Disease Prism Reagent Qualfication P1 | Prism Abbott Release Positive Channel A Result Q6.P1 | Standard Acceptance Criteria Prism Abbott Release Positive Channel A Result Q6.P1 Equal to Positive
Infectious Disease PRISM Reagent Qualfication | Infectious Disease Prism Reagent Qualfication P1 | Prism Abbott Release Positive Channel B Result Q7.P1 | Standard Acceptance Criteria Prism Abbott Release Positive Channel B Result Q7.P1 Equal to Positive
Infectious Disease PRISM Reagent Qualfication | Infectious Disease Prism Reagent Qualfication P1 | Prism Calibrators and External Controls Acceptable? Q8.P1 | Standard Acceptance Criteria Prism Calibrators and External Controls Acceptable? Q8.P1 Equal to Acceptable
Infectious Disease PRISM Reagent Qualfication | Infectious Disease Prism Reagent Qualfication P1 | Problem with Equipment Q9.P1 | Standard Acceptance Criteria Problem with Equipment Q9.P1 Equal to No
To be removed | Capture-P Washstation Daily P1 | Empty Waste and Clean Q1.P1 | Standard Acceptance Criteria Empty Waste and Clean Q1.P1 Equal to Performed - OK
To be removed | Capture-P Washstation Daily P1 | Check/fill PBS & DI Water Q2.P1 | Standard Acceptance Criteria Check/fill PBS & DI Water Q2.P1 Equal to Performed - OK
To be removed | Capture-P Washstation Daily P1 | Prime / Check Probe Q3.P1 | Standard Acceptance Criteria Prime / Check Probe Q3.P1 Equal to Performed - OK
To be removed | Capture-P Washstation Daily P1 | Residual Saline Verification Q4.P1 | Standard Acceptance Criteria Residual Saline Verification Q4.P1 Equal to Performed - OK
To be removed | Capture-P Washstation Daily P1 | Record Heatblock Incubator Temperature Q5.P1 | Standard Acceptance Criteria Record Heatblock Incubator Temperature Q5.P1 Between 36 to 38
To be removed | Capture-P Washstation Daily PBS P1.P2(Con) | Buffered Saline PH Result Q1.P1.P2(Con) | Standard Acceptance Criteria Buffered Saline PH Result Q1.P1.P2(Con) Between 6.8 to 7.4
To be removed | Capture-P Washstation Daily PBS P1.P2(Con) | New PBS Expiration Date Q2.P1.P2(Con) | Standard Acceptance Criteria New PBS Expiration Date Q2.P1.P2(Con) Greater than or equal to Today
Additional Vacutainer Tubes Used | Problem with QC P1.P2(Con) | Problems with QC Q1.P1.P2(Con) | Standard Acceptance Criteria Problems with QC Q1.P1.P2(Con) Equal to No
Coulter AcT 2 Diff Shutdown Daily | Shutdown P4 | Clean and Shut Down Instrument? Q1.P4 | Standard Acceptance Criteria Clean and Shut Down Instrument? Q1.P4 Equal to Performed - OK
Coulter AcT 2 Diff Shutdown Daily | Shutdown P4 | Problem with Equipment Q2.P4 | Standard Acceptance Criteria Problem with Equipment Q2.P4 Equal to No
Equipment Problem Log - Copy-RLP | 11-Re-activate Out of Service Equipment P1.P2(Con).P2(Con).P2(Con) | Re-activate Out of Service Equipment Q1.P1.P2(Con).P2(Con).P2(Con) | Standard Acceptance Criteria Re-activate Out of Service Equipment Q1.P1.P2(Con).P2(Con).P2(Con) Equal to Instructions Acknowledged
Equipment Problem Log - Copy-RLP | 11-Re-activate Out of Service Equipment P1.P2(Con).P2(Con).P3(Con).P2(Con).P2(Con) | Re-activate Out of Service Equipment Q1.P1.P2(Con).P2(Con).P3(Con).P2(Con).P2(Con) | Standard Acceptance Criteria Re-activate Out of Service Equipment Q1.P1.P2(Con).P2(Con).P3(Con).P2(Con).P2(Con) Equal to Instructions Acknowledged
Equipment Problem Log - Copy-RLP | 11-Re-activate Out of Service Equipment P1.P2(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P2(Con) | Re-activate Out of Service Equipment Q1.P1.P2(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P2(Con) | Standard Acceptance Criteria Re-activate Out of Service Equipment Q1.P1.P2(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P2(Con) Equal to Instructions Acknowledged
Equipment Problem Log - Copy-RLP | 11-Re-activate Out of Service Equipment P1.P2(Con).P4(Con).P2(Con).P2(Con).P2(Con).P4(Con).P2(Con) | Re-activate Out of Service Equipment Q1.P1.P2(Con).P4(Con).P2(Con).P2(Con).P2(Con).P4(Con).P2(Con) | Standard Acceptance Criteria Re-activate Out of Service Equipment Q1.P1.P2(Con).P4(Con).P2(Con).P2(Con).P2(Con).P4(Con).P2(Con) Equal to Instructions Acknowledged
Equipment Problem Log - Copy-RLP | 11-Re-activate Out of Service Equipment P1.P2(Con).P4(Con).P2(Con).P2(Con).P2(Con).P4(Con).P3(Con).P2(Con).P2(Con) | Re-activate Out of Service Equipment Q1.P1.P2(Con).P4(Con).P2(Con).P2(Con).P2(Con).P4(Con).P3(Con).P2(Con).P2(Con) | Standard Acceptance Criteria Re-activate Out of Service Equipment Q1.P1.P2(Con).P4(Con).P2(Con).P2(Con).P2(Con).P4(Con).P3(Con).P2(Con).P2(Con) Equal to Instructions Acknowledged
Equipment Problem Log - Copy-RLP | 11-Re-activate Out of Service Equipment P1.P2(Con).P4(Con).P2(Con).P2(Con).P2(Con).P5(Con).P2(Con).P2(Con) | Re-activate Out of Service Equipment Q1.P1.P2(Con).P4(Con).P2(Con).P2(Con).P2(Con).P5(Con).P2(Con).P2(Con) | Standard Acceptance Criteria Re-activate Out of Service Equipment Q1.P1.P2(Con).P4(Con).P2(Con).P2(Con).P2(Con).P5(Con).P2(Con).P2(Con) Equal to Instructions Acknowledged
Equipment Problem Log - Copy-RLP | 11-Re-activate Out of Service Equipment P1.P2(Con).P4(Con).P2(Con).P2(Con).P2(Con).P5(Con).P2(Con).P3(Con).P2(Con).P2(Con) | Re-activate Out of Service Equipment Q1.P1.P2(Con).P4(Con).P2(Con).P2(Con).P2(Con).P5(Con).P2(Con).P3(Con).P2(Con).P2(Con) | Standard Acceptance Criteria Re-activate Out of Service Equipment Q1.P1.P2(Con).P4(Con).P2(Con).P2(Con).P2(Con).P5(Con).P2(Con).P3(Con).P2(Con).P2(Con) Equal to Instructions Acknowledged
Equipment Problem Log - Copy-RLP | 11-Re-activate Out of Service Equipment P1.P2(Con).P4(Con).P2(Con).P2(Con).P2(Con).P5(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con) | Re-activate Out of Service Equipment Q1.P1.P2(Con).P4(Con).P2(Con).P2(Con).P2(Con).P5(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con) | Standard Acceptance Criteria Re-activate Out of Service Equipment Q1.P1.P2(Con).P4(Con).P2(Con).P2(Con).P2(Con).P5(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con) Equal to Instructions Acknowledged
Equipment Problem Log - Copy-RLP | 11-Re-activate Out of Service Equipment P1.P2(Con).P4(Con).P2(Con).P2(Con).P3(Con).P2(Con) | Re-activate Out of Service Equipment Q1.P1.P2(Con).P4(Con).P2(Con).P2(Con).P3(Con).P2(Con) | Standard Acceptance Criteria Re-activate Out of Service Equipment Q1.P1.P2(Con).P4(Con).P2(Con).P2(Con).P3(Con).P2(Con) Equal to Instructions Acknowledged
Equipment Problem Log - Copy-RLP | 11-Re-activate Out of Service Equipment P1.P2(Con).P4(Con).P2(Con).P2(Con).P3(Con).P3(Con).P2(Con).P2(Con) | Re-activate Out of Service Equipment Q1.P1.P2(Con).P4(Con).P2(Con).P2(Con).P3(Con).P3(Con).P2(Con).P2(Con) | Standard Acceptance Criteria Re-activate Out of Service Equipment Q1.P1.P2(Con).P4(Con).P2(Con).P2(Con).P3(Con).P3(Con).P2(Con).P2(Con) Equal to Instructions Acknowledged
Equipment Problem Log - Copy-RLP | 11-Re-activate Out of Service Equipment P1.P2(Con).P4(Con).P3(Con).P2(Con).P2(Con).P2(Con) | Re-activate Out of Service Equipment Q1.P1.P2(Con).P4(Con).P3(Con).P2(Con).P2(Con).P2(Con) | Standard Acceptance Criteria Re-activate Out of Service Equipment Q1.P1.P2(Con).P4(Con).P3(Con).P2(Con).P2(Con).P2(Con) Equal to Instructions Acknowledged
Equipment Problem Log - Copy-RLP | 11-Re-activate Out of Service Equipment P1.P2(Con).P4(Con).P3(Con).P2(Con).P2(Con).P3(Con).P2(Con).P2(Con) | Re-activate Out of Service Equipment Q1.P1.P2(Con).P4(Con).P3(Con).P2(Con).P2(Con).P3(Con).P2(Con).P2(Con) | Standard Acceptance Criteria Re-activate Out of Service Equipment Q1.P1.P2(Con).P4(Con).P3(Con).P2(Con).P2(Con).P3(Con).P2(Con).P2(Con) Equal to Instructions Acknowledged
Equipment Problem Log - Copy-RLP | 11-Re-activate Out of Service Equipment P1.P2(Con).P4(Con).P3(Con).P2(Con).P3(Con).P2(Con).P2(Con) | Re-activate Out of Service Equipment Q1.P1.P2(Con).P4(Con).P3(Con).P2(Con).P3(Con).P2(Con).P2(Con) | Standard Acceptance Criteria Re-activate Out of Service Equipment Q1.P1.P2(Con).P4(Con).P3(Con).P2(Con).P3(Con).P2(Con).P2(Con) Equal to Instructions Acknowledged
Equipment Problem Log - Copy-RLP | 11-Re-activate Out of Service Equipment P1.P2(Con).P4(Con).P3(Con).P3(Con).P2(Con) | Re-activate Out of Service Equipment Q1.P1.P2(Con).P4(Con).P3(Con).P3(Con).P2(Con) | Standard Acceptance Criteria Re-activate Out of Service Equipment Q1.P1.P2(Con).P4(Con).P3(Con).P3(Con).P2(Con) Equal to Instructions Acknowledged
Equipment Problem Log - Copy-RLP | 11-Re-activate Out of Service Equipment P1.P2(Con).P4(Con).P3(Con).P3(Con).P3(Con).P2(Con).P2(Con) | Re-activate Out of Service Equipment Q1.P1.P2(Con).P4(Con).P3(Con).P3(Con).P3(Con).P2(Con).P2(Con) | Standard Acceptance Criteria Re-activate Out of Service Equipment Q1.P1.P2(Con).P4(Con).P3(Con).P3(Con).P3(Con).P2(Con).P2(Con) Equal to Instructions Acknowledged
Equipment Problem Log - Copy-RLP | 11-Re-activate Out of Service Equipment P1.P3(Con).P2(Con).P2(Con).P2(Con).P2(Con).P2(Con) | Re-activate Out of Service Equipment Q1.P1.P3(Con).P2(Con).P2(Con).P2(Con).P2(Con).P2(Con) | Standard Acceptance Criteria Re-activate Out of Service Equipment Q1.P1.P3(Con).P2(Con).P2(Con).P2(Con).P2(Con).P2(Con) Equal to Instructions Acknowledged
Equipment Problem Log - Copy-RLP | 11-Re-activate Out of Service Equipment P1.P3(Con).P2(Con).P2(Con).P2(Con).P2(Con).P3(Con).P2(Con).P2(Con) | Re-activate Out of Service Equipment Q1.P1.P3(Con).P2(Con).P2(Con).P2(Con).P2(Con).P3(Con).P2(Con).P2(Con) | Standard Acceptance Criteria Re-activate Out of Service Equipment Q1.P1.P3(Con).P2(Con).P2(Con).P2(Con).P2(Con).P3(Con).P2(Con).P2(Con) Equal to Instructions Acknowledged
Equipment Problem Log - Copy-RLP | 11-Re-activate Out of Service Equipment P1.P3(Con).P2(Con).P2(Con).P2(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P2(Con) | Re-activate Out of Service Equipment Q1.P1.P3(Con).P2(Con).P2(Con).P2(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P2(Con) | Standard Acceptance Criteria Re-activate Out of Service Equipment Q1.P1.P3(Con).P2(Con).P2(Con).P2(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P2(Con) Equal to Instructions Acknowledged
Equipment Problem Log - Copy-RLP | 11-Re-activate Out of Service Equipment P1.P3(Con).P2(Con).P2(Con).P2(Con).P3(Con).P2(Con).P2(Con) | Re-activate Out of Service Equipment Q1.P1.P3(Con).P2(Con).P2(Con).P2(Con).P3(Con).P2(Con).P2(Con) | Standard Acceptance Criteria Re-activate Out of Service Equipment Q1.P1.P3(Con).P2(Con).P2(Con).P2(Con).P3(Con).P2(Con).P2(Con) Equal to Instructions Acknowledged
Equipment Problem Log - Copy-RLP | 11-Re-activate Out of Service Equipment P1.P3(Con).P2(Con).P2(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P2(Con) | Re-activate Out of Service Equipment Q1.P1.P3(Con).P2(Con).P2(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P2(Con) | Standard Acceptance Criteria Re-activate Out of Service Equipment Q1.P1.P3(Con).P2(Con).P2(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P2(Con) Equal to Instructions Acknowledged
Equipment Problem Log - Copy-RLP | 11-Re-activate Out of Service Equipment P1.P3(Con).P2(Con).P2(Con).P3(Con).P2(Con) | Re-activate Out of Service Equipment Q1.P1.P3(Con).P2(Con).P2(Con).P3(Con).P2(Con) | Standard Acceptance Criteria Re-activate Out of Service Equipment Q1.P1.P3(Con).P2(Con).P2(Con).P3(Con).P2(Con) Equal to Instructions Acknowledged
Equipment Problem Log - Copy-RLP | 11-Re-activate Out of Service Equipment P1.P3(Con).P2(Con).P2(Con).P3(Con).P3(Con).P2(Con).P2(Con) | Re-activate Out of Service Equipment Q1.P1.P3(Con).P2(Con).P2(Con).P3(Con).P3(Con).P2(Con).P2(Con) | Standard Acceptance Criteria Re-activate Out of Service Equipment Q1.P1.P3(Con).P2(Con).P2(Con).P3(Con).P3(Con).P2(Con).P2(Con) Equal to Instructions Acknowledged
Equipment Problem Log - Copy-RLP | 11-Re-activate Out of Service Equipment P1.P3(Con).P2(Con).P2(Con).P3(Con).P3(Con).P2(Con).P3(Con).P2(Con).P2(Con) | Re-activate Out of Service Equipment Q1.P1.P3(Con).P2(Con).P2(Con).P3(Con).P3(Con).P2(Con).P3(Con).P2(Con).P2(Con) | Standard Acceptance Criteria Re-activate Out of Service Equipment Q1.P1.P3(Con).P2(Con).P2(Con).P3(Con).P3(Con).P2(Con).P3(Con).P2(Con).P2(Con) Equal to Instructions Acknowledged
Equipment Problem Log - Copy-RLP | 11-Re-activate Out of Service Equipment P1.P3(Con).P2(Con).P2(Con).P3(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P2(Con) | Re-activate Out of Service Equipment Q1.P1.P3(Con).P2(Con).P2(Con).P3(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P2(Con) | Standard Acceptance Criteria Re-activate Out of Service Equipment Q1.P1.P3(Con).P2(Con).P2(Con).P3(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P2(Con) Equal to Instructions Acknowledged
Equipment Problem Log - Copy-RLP | 11-Re-activate Out of Service Equipment P1.P3(Con).P2(Con).P2(Con).P3(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P2(Con) | Re-activate Out of Service Equipment Q1.P1.P3(Con).P2(Con).P2(Con).P3(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P2(Con) | Standard Acceptance Criteria Re-activate Out of Service Equipment Q1.P1.P3(Con).P2(Con).P2(Con).P3(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P2(Con) Equal to Instructions Acknowledged
Equipment Problem Log - Copy-RLP | 11-Re-activate Out of Service Equipment P1.P3(Con).P2(Con).P2(Con).P3(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P2(Con) | Re-activate Out of Service Equipment Q1.P1.P3(Con).P2(Con).P2(Con).P3(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P2(Con) | Standard Acceptance Criteria Re-activate Out of Service Equipment Q1.P1.P3(Con).P2(Con).P2(Con).P3(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P2(Con) Equal to Instructions Acknowledged
Equipment Problem Log - Copy-RLP | 11-Re-activate Out of Service Equipment P1.P3(Con).P2(Con).P2(Con).P3(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P2(Con) | Re-activate Out of Service Equipment Q1.P1.P3(Con).P2(Con).P2(Con).P3(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P2(Con) | Standard Acceptance Criteria Re-activate Out of Service Equipment Q1.P1.P3(Con).P2(Con).P2(Con).P3(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P2(Con) Equal to Instructions Acknowledged
Equipment Problem Log - Copy-RLP | 11-Re-activate Out of Service Equipment P1.P3(Con).P2(Con).P2(Con).P3(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P2(Con) | Re-activate Out of Service Equipment Q1.P1.P3(Con).P2(Con).P2(Con).P3(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P2(Con) | Standard Acceptance Criteria Re-activate Out of Service Equipment Q1.P1.P3(Con).P2(Con).P2(Con).P3(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P2(Con) Equal to Instructions Acknowledged
Equipment Problem Log - Copy-RLP | 11-Re-activate Out of Service Equipment P1.P3(Con).P2(Con).P2(Con).P3(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con) | Re-activate Out of Service Equipment Q1.P1.P3(Con).P2(Con).P2(Con).P3(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con) | Standard Acceptance Criteria Re-activate Out of Service Equipment Q1.P1.P3(Con).P2(Con).P2(Con).P3(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con) Equal to Instructions Acknowledged
Equipment Problem Log - Copy-RLP | 11-Re-activate Out of Service Equipment P1.P3(Con).P3(Con).P2(Con).P2(Con).P2(Con) | Re-activate Out of Service Equipment Q1.P1.P3(Con).P3(Con).P2(Con).P2(Con).P2(Con) | Standard Acceptance Criteria Re-activate Out of Service Equipment Q1.P1.P3(Con).P3(Con).P2(Con).P2(Con).P2(Con) Equal to Instructions Acknowledged
Equipment Problem Log - Copy-RLP | 11-Re-activate Out of Service Equipment P1.P3(Con).P3(Con).P2(Con).P2(Con).P3(Con).P2(Con).P2(Con) | Re-activate Out of Service Equipment Q1.P1.P3(Con).P3(Con).P2(Con).P2(Con).P3(Con).P2(Con).P2(Con) | Standard Acceptance Criteria Re-activate Out of Service Equipment Q1.P1.P3(Con).P3(Con).P2(Con).P2(Con).P3(Con).P2(Con).P2(Con) Equal to Instructions Acknowledged
Equipment Problem Log - Copy-RLP | 11-Re-activate Out of Service Equipment P1.P3(Con).P3(Con).P2(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P2(Con) | Re-activate Out of Service Equipment Q1.P1.P3(Con).P3(Con).P2(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P2(Con) | Standard Acceptance Criteria Re-activate Out of Service Equipment Q1.P1.P3(Con).P3(Con).P2(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P2(Con) Equal to Instructions Acknowledged
Equipment Problem Log - Copy-RLP | 11-Re-activate Out of Service Equipment P1.P3(Con).P3(Con).P2(Con).P3(Con).P2(Con).P2(Con) | Re-activate Out of Service Equipment Q1.P1.P3(Con).P3(Con).P2(Con).P3(Con).P2(Con).P2(Con) | Standard Acceptance Criteria Re-activate Out of Service Equipment Q1.P1.P3(Con).P3(Con).P2(Con).P3(Con).P2(Con).P2(Con) Equal to Instructions Acknowledged
Equipment Problem Log - Copy-RLP | 11-Re-activate Out of Service Equipment P1.P3(Con).P3(Con).P3(Con).P2(Con) | Re-activate Out of Service Equipment Q1.P1.P3(Con).P3(Con).P3(Con).P2(Con) | Standard Acceptance Criteria Re-activate Out of Service Equipment Q1.P1.P3(Con).P3(Con).P3(Con).P2(Con) Equal to Instructions Acknowledged
Equipment Problem Log - Copy-RLP | 11-Re-activate Out of Service Equipment P1.P3(Con).P3(Con).P3(Con).P3(Con).P2(Con).P2(Con) | Re-activate Out of Service Equipment Q1.P1.P3(Con).P3(Con).P3(Con).P3(Con).P2(Con).P2(Con) | Standard Acceptance Criteria Re-activate Out of Service Equipment Q1.P1.P3(Con).P3(Con).P3(Con).P3(Con).P2(Con).P2(Con) Equal to Instructions Acknowledged
Equipment Problem Log - Copy0316 | 4-Service Call P1.P2(Con).P4(Con).P2(Con) | Problem Resolved? Q3.P1.P2(Con).P4(Con).P2(Con) | Standard Acceptance Criteria Problem Resolved? Q3.P1.P2(Con).P4(Con).P2(Con) Equal to Yes
Equipment Problem Log - Copy0316 | 6-Qualifications/Validation P1.P2(Con).P4(Con).P2(Con).P2(Con).P2(Con) | Qualification/Validation Acceptable? Q3.P1.P2(Con).P4(Con).P2(Con).P2(Con).P2(Con) | Standard Acceptance Criteria Qualification/Validation Acceptable? Q3.P1.P2(Con).P4(Con).P2(Con).P2(Con).P2(Con) Equal to Yes
Equipment Problem Log - Copy0316 | 11-Re-activate Out of Service Equipment P1.P2(Con).P4(Con).P2(Con).P2(Con).P2(Con).P4(Con).P2(Con) | Re-activate Out of Service Equipment Q1.P1.P2(Con).P4(Con).P2(Con).P2(Con).P2(Con).P4(Con).P2(Con) | Standard Acceptance Criteria Re-activate Out of Service Equipment Q1.P1.P2(Con).P4(Con).P2(Con).P2(Con).P2(Con).P4(Con).P2(Con) Equal to Instructions Acknowledged
Equipment Problem Log - Copy0316 | 11-Re-activate Out of Service Equipment P1.P2(Con).P4(Con).P2(Con).P2(Con).P2(Con).P4(Con).P3(Con).P2(Con).P2(Con) | Re-activate Out of Service Equipment Q1.P1.P2(Con).P4(Con).P2(Con).P2(Con).P2(Con).P4(Con).P3(Con).P2(Con).P2(Con) | Standard Acceptance Criteria Re-activate Out of Service Equipment Q1.P1.P2(Con).P4(Con).P2(Con).P2(Con).P2(Con).P4(Con).P3(Con).P2(Con).P2(Con) Equal to Instructions Acknowledged
Equipment Problem Log - Copy0316 | 11-Re-activate Out of Service Equipment P1.P2(Con).P4(Con).P2(Con).P2(Con).P2(Con).P5(Con).P2(Con).P2(Con) | Re-activate Out of Service Equipment Q1.P1.P2(Con).P4(Con).P2(Con).P2(Con).P2(Con).P5(Con).P2(Con).P2(Con) | Standard Acceptance Criteria Re-activate Out of Service Equipment Q1.P1.P2(Con).P4(Con).P2(Con).P2(Con).P2(Con).P5(Con).P2(Con).P2(Con) Equal to Instructions Acknowledged
Equipment Problem Log - Copy0316 | 11-Re-activate Out of Service Equipment P1.P2(Con).P4(Con).P2(Con).P2(Con).P2(Con).P5(Con).P2(Con).P3(Con).P2(Con).P2(Con) | Re-activate Out of Service Equipment Q1.P1.P2(Con).P4(Con).P2(Con).P2(Con).P2(Con).P5(Con).P2(Con).P3(Con).P2(Con).P2(Con) | Standard Acceptance Criteria Re-activate Out of Service Equipment Q1.P1.P2(Con).P4(Con).P2(Con).P2(Con).P2(Con).P5(Con).P2(Con).P3(Con).P2(Con).P2(Con) Equal to Instructions Acknowledged
Equipment Problem Log - Copy0316 | 11-Re-activate Out of Service Equipment P1.P2(Con).P4(Con).P2(Con).P2(Con).P2(Con).P5(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con) | Re-activate Out of Service Equipment Q1.P1.P2(Con).P4(Con).P2(Con).P2(Con).P2(Con).P5(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con) | Standard Acceptance Criteria Re-activate Out of Service Equipment Q1.P1.P2(Con).P4(Con).P2(Con).P2(Con).P2(Con).P5(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con) Equal to Instructions Acknowledged
Equipment Problem Log - Copy0316 | 11-Re-activate Out of Service Equipment P1.P2(Con).P4(Con).P2(Con).P2(Con).P3(Con).P2(Con) | Re-activate Out of Service Equipment Q1.P1.P2(Con).P4(Con).P2(Con).P2(Con).P3(Con).P2(Con) | Standard Acceptance Criteria Re-activate Out of Service Equipment Q1.P1.P2(Con).P4(Con).P2(Con).P2(Con).P3(Con).P2(Con) Equal to Instructions Acknowledged
Equipment Problem Log - Copy0316 | 11-Re-activate Out of Service Equipment P1.P2(Con).P4(Con).P2(Con).P2(Con).P3(Con).P3(Con).P2(Con).P2(Con) | Re-activate Out of Service Equipment Q1.P1.P2(Con).P4(Con).P2(Con).P2(Con).P3(Con).P3(Con).P2(Con).P2(Con) | Standard Acceptance Criteria Re-activate Out of Service Equipment Q1.P1.P2(Con).P4(Con).P2(Con).P2(Con).P3(Con).P3(Con).P2(Con).P2(Con) Equal to Instructions Acknowledged
Equipment Problem Log - Copy0316 | 4-Service Call P1.P3(Con).P2(Con) | Problem Resolved? Q3.P1.P3(Con).P2(Con) | Standard Acceptance Criteria Problem Resolved? Q3.P1.P3(Con).P2(Con) Equal to Yes
Equipment Problem Log - Copy0316 | 6-Qualifications/Validation P1.P3(Con).P2(Con).P2(Con).P2(Con) | Qualification/Validation Acceptable? Q3.P1.P3(Con).P2(Con).P2(Con).P2(Con) | Standard Acceptance Criteria Qualification/Validation Acceptable? Q3.P1.P3(Con).P2(Con).P2(Con).P2(Con) Equal to Yes
Equipment Problem Log - Copy0316 | 11-Re-activate Out of Service Equipment P1.P3(Con).P2(Con).P2(Con).P2(Con).P2(Con).P2(Con) | Re-activate Out of Service Equipment Q1.P1.P3(Con).P2(Con).P2(Con).P2(Con).P2(Con).P2(Con) | Standard Acceptance Criteria Re-activate Out of Service Equipment Q1.P1.P3(Con).P2(Con).P2(Con).P2(Con).P2(Con).P2(Con) Equal to Instructions Acknowledged
Equipment Problem Log - Copy0316 | 11-Re-activate Out of Service Equipment P1.P3(Con).P2(Con).P2(Con).P2(Con).P2(Con).P3(Con).P2(Con).P2(Con) | Re-activate Out of Service Equipment Q1.P1.P3(Con).P2(Con).P2(Con).P2(Con).P2(Con).P3(Con).P2(Con).P2(Con) | Standard Acceptance Criteria Re-activate Out of Service Equipment Q1.P1.P3(Con).P2(Con).P2(Con).P2(Con).P2(Con).P3(Con).P2(Con).P2(Con) Equal to Instructions Acknowledged
Equipment Problem Log - Copy0316 | 11-Re-activate Out of Service Equipment P1.P3(Con).P2(Con).P2(Con).P2(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P2(Con) | Re-activate Out of Service Equipment Q1.P1.P3(Con).P2(Con).P2(Con).P2(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P2(Con) | Standard Acceptance Criteria Re-activate Out of Service Equipment Q1.P1.P3(Con).P2(Con).P2(Con).P2(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P2(Con) Equal to Instructions Acknowledged
Equipment Problem Log - Copy0316 | 11-Re-activate Out of Service Equipment P1.P3(Con).P2(Con).P2(Con).P2(Con).P3(Con).P2(Con).P2(Con) | Re-activate Out of Service Equipment Q1.P1.P3(Con).P2(Con).P2(Con).P2(Con).P3(Con).P2(Con).P2(Con) | Standard Acceptance Criteria Re-activate Out of Service Equipment Q1.P1.P3(Con).P2(Con).P2(Con).P2(Con).P3(Con).P2(Con).P2(Con) Equal to Instructions Acknowledged
Equipment Problem Log - Copy0316 | 11-Re-activate Out of Service Equipment P1.P3(Con).P2(Con).P2(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P2(Con) | Re-activate Out of Service Equipment Q1.P1.P3(Con).P2(Con).P2(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P2(Con) | Standard Acceptance Criteria Re-activate Out of Service Equipment Q1.P1.P3(Con).P2(Con).P2(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P2(Con) Equal to Instructions Acknowledged
Equipment Problem Log - Copy0316 | 11-Re-activate Out of Service Equipment P1.P3(Con).P2(Con).P2(Con).P3(Con).P2(Con) | Re-activate Out of Service Equipment Q1.P1.P3(Con).P2(Con).P2(Con).P3(Con).P2(Con) | Standard Acceptance Criteria Re-activate Out of Service Equipment Q1.P1.P3(Con).P2(Con).P2(Con).P3(Con).P2(Con) Equal to Instructions Acknowledged
Equipment Problem Log - Copy0316 | 11-Re-activate Out of Service Equipment P1.P3(Con).P2(Con).P2(Con).P3(Con).P3(Con).P2(Con).P2(Con) | Re-activate Out of Service Equipment Q1.P1.P3(Con).P2(Con).P2(Con).P3(Con).P3(Con).P2(Con).P2(Con) | Standard Acceptance Criteria Re-activate Out of Service Equipment Q1.P1.P3(Con).P2(Con).P2(Con).P3(Con).P3(Con).P2(Con).P2(Con) Equal to Instructions Acknowledged
Equipment Problem Log - Copy0316 | 11-Re-activate Out of Service Equipment P1.P3(Con).P2(Con).P2(Con).P3(Con).P3(Con).P2(Con).P3(Con).P2(Con).P2(Con) | Re-activate Out of Service Equipment Q1.P1.P3(Con).P2(Con).P2(Con).P3(Con).P3(Con).P2(Con).P3(Con).P2(Con).P2(Con) | Standard Acceptance Criteria Re-activate Out of Service Equipment Q1.P1.P3(Con).P2(Con).P2(Con).P3(Con).P3(Con).P2(Con).P3(Con).P2(Con).P2(Con) Equal to Instructions Acknowledged
Equipment Problem Log - Copy0316 | 11-Re-activate Out of Service Equipment P1.P3(Con).P2(Con).P2(Con).P3(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P2(Con) | Re-activate Out of Service Equipment Q1.P1.P3(Con).P2(Con).P2(Con).P3(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P2(Con) | Standard Acceptance Criteria Re-activate Out of Service Equipment Q1.P1.P3(Con).P2(Con).P2(Con).P3(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P2(Con) Equal to Instructions Acknowledged
Equipment Problem Log - Copy0316 | 11-Re-activate Out of Service Equipment P1.P3(Con).P2(Con).P2(Con).P3(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P2(Con) | Re-activate Out of Service Equipment Q1.P1.P3(Con).P2(Con).P2(Con).P3(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P2(Con) | Standard Acceptance Criteria Re-activate Out of Service Equipment Q1.P1.P3(Con).P2(Con).P2(Con).P3(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P2(Con) Equal to Instructions Acknowledged
Equipment Problem Log - Copy0316 | 11-Re-activate Out of Service Equipment P1.P3(Con).P2(Con).P2(Con).P3(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P2(Con) | Re-activate Out of Service Equipment Q1.P1.P3(Con).P2(Con).P2(Con).P3(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P2(Con) | Standard Acceptance Criteria Re-activate Out of Service Equipment Q1.P1.P3(Con).P2(Con).P2(Con).P3(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P2(Con) Equal to Instructions Acknowledged
Equipment Problem Log - Copy0316 | 11-Re-activate Out of Service Equipment P1.P3(Con).P2(Con).P2(Con).P3(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P2(Con) | Re-activate Out of Service Equipment Q1.P1.P3(Con).P2(Con).P2(Con).P3(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P2(Con) | Standard Acceptance Criteria Re-activate Out of Service Equipment Q1.P1.P3(Con).P2(Con).P2(Con).P3(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P2(Con) Equal to Instructions Acknowledged
Equipment Problem Log - Copy0316 | 11-Re-activate Out of Service Equipment P1.P3(Con).P2(Con).P2(Con).P3(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P2(Con) | Re-activate Out of Service Equipment Q1.P1.P3(Con).P2(Con).P2(Con).P3(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P2(Con) | Standard Acceptance Criteria Re-activate Out of Service Equipment Q1.P1.P3(Con).P2(Con).P2(Con).P3(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P2(Con) Equal to Instructions Acknowledged
Equipment Problem Log - Copy0316 | 11-Re-activate Out of Service Equipment P1.P3(Con).P2(Con).P2(Con).P3(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con) | Re-activate Out of Service Equipment Q1.P1.P3(Con).P2(Con).P2(Con).P3(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con) | Standard Acceptance Criteria Re-activate Out of Service Equipment Q1.P1.P3(Con).P2(Con).P2(Con).P3(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con).P3(Con).P2(Con) Equal to Instructions Acknowledged
WNV Cutoff/Correlation As Needed | WNV Cutoff/Correlation P1 | WNV Negative External Control Day 1 Valid Q14.P1 | Standard Acceptance Criteria WNV Negative External Control Day 1 Valid Q14.P1 Equal to Yes
WNV Cutoff/Correlation As Needed | WNV Cutoff/Correlation P1 | WNV Positive External Control Day 1 Valid Q15.P1 | Standard Acceptance Criteria WNV Positive External Control Day 1 Valid Q15.P1 Equal to Yes
WNV Cutoff/Correlation As Needed | WNV Cutoff/Correlation P1 | Non-Reactive Day 1 Valid Q7.P1 | Standard Acceptance Criteria Non-Reactive Day 1 Valid Q7.P1 Equal to Yes
WNV Cutoff/Correlation As Needed | WNV Cutoff/Correlation P1 | Reactive Day 1 Valid Q8.P1 | Standard Acceptance Criteria Reactive Day 1 Valid Q8.P1 Equal to Yes
WNV Cutoff/Correlation As Needed | WNV Cutoff/Correlation P1 | WNV Negative External Control Day 2 Valid Q16.P1 | Standard Acceptance Criteria WNV Negative External Control Day 2 Valid Q16.P1 Equal to Yes
WNV Cutoff/Correlation As Needed | WNV Cutoff/Correlation P1 | WNV Positive External Control Day 2 Valid Q17.P1 | Standard Acceptance Criteria WNV Positive External Control Day 2 Valid Q17.P1 Equal to Yes
WNV Cutoff/Correlation As Needed | WNV Cutoff/Correlation P1 | Non-Reactive Day 2 Valid Q21.P1 | Standard Acceptance Criteria Non-Reactive Day 2 Valid Q21.P1 Equal to Yes
WNV Cutoff/Correlation As Needed | WNV Cutoff/Correlation P1 | Reactive Day 2 Valid Q22.P1 | Standard Acceptance Criteria Reactive Day 2 Valid Q22.P1 Equal to Yes
WNV Cutoff/Correlation As Needed | WNV Cutoff/Correlation P1 | Is Non-Reactive Analyte RLU on Day 1 Close to Non-Reactive Analyte on Day 2 (are Both Non-reactive?) Q9.P1 | Standard Acceptance Criteria Is Non-Reactive Analyte RLU on Day 1 Close to Non-Reactive Analyte on Day 2 (are Both Non-reactive?) Q9.P1 Equal to Yes
WNV Cutoff/Correlation As Needed | WNV Cutoff/Correlation P1 | Is Reactive Analyte RLU on Day 1 Close to Reactive Analyte RLU on Day 2 (are Both Reactive?) Q10.P1 | Standard Acceptance Criteria Is Reactive Analyte RLU on Day 1 Close to Reactive Analyte RLU on Day 2 (are Both Reactive?) Q10.P1 Equal to Yes
WNV Cutoff/Correlation As Needed | WNV Cutoff/Correlation P1 | Are the Runs Acceptable? (All Controls/calibrators Results Are as Expected?) Q11.P1 | Standard Acceptance Criteria Are the Runs Acceptable? (All Controls/calibrators Results Are as Expected?) Q11.P1 Equal to Acceptable
WNV Cutoff/Correlation As Needed | WNV Cutoff/Correlation P1 | Is Correlation of Instruments Acceptable? (Answer Only If 2 Instruments Are Used) Q12.P1 | Standard Acceptance Criteria Is Correlation of Instruments Acceptable? (Answer Only If 2 Instruments Are Used) Q12.P1 Equal to Acceptable
WNV Cutoff/Correlation As Needed | WNV Cutoff/Correlation P1 | Problem with Equipment Q13.P1 | Standard Acceptance Criteria Problem with Equipment Q13.P1 Equal to No
Bioarray PCR Thermal Cycler Weekly | PCR Thermal Cycler Weekly P2 | Clean Thermal Cycler Q1.P2 | Standard Acceptance Criteria Clean Thermal Cycler Q1.P2 Equal to Performed - OK
Bioarray Hermle Centrifuge Monthly | BioArray Hermle Monthly P2 | Clean Hermle Centrifuge and Lubricate Rotors Q1.P2 | Standard Acceptance Criteria Clean Hermle Centrifuge and Lubricate Rotors Q1.P2 Equal to Performed - OK
BioArray PCR Hood Area Wipe Test Monthly | BioArray PCR Hood Area Wipe Test Monthly P2 | Final Result of Wipe Test for PCR Hood Area? Q1.P2 | Standard Acceptance Criteria Final Result of Wipe Test for PCR Hood Area? Q1.P2 Equal to NTD (No Type Determined) - Acceptable
BioArray DNA Addition Area Wipe Test Monthly | BioArray DNA Addition Area Wipe Test Monthly P3 | Final Result of Wipe Test for DNA Addition Area? Q1.P3 | Standard Acceptance Criteria Final Result of Wipe Test for DNA Addition Area? Q1.P3 Equal to NTD (No Type Determined) - Acceptable
BioArray Post PCR Processing Area Wipe Test Mnthly | BioArray Post PCR Processing Area Wipe Test Monthly P3 | Final Result of Wipe Test for Post PCR Processing Area? Q1.P3 | Standard Acceptance Criteria Final Result of Wipe Test for Post PCR Processing Area? Q1.P3 Equal to NTD (No Type Determined) - Acceptable
Neo Correlation FWD ABORH Testing | FWD ABORH (Automated) P1 | Results Acceptable? Q13.P1 | Standard Acceptance Criteria Results Acceptable? Q13.P1 Equal to Acceptable
Neo Correlation FWD ABORH Testing | FWD ABORH (Automated) P1 | Problem with Equipment Q14.P1 | Standard Acceptance Criteria Problem with Equipment Q14.P1 Equal to No
RPR Daily | RPR Daily P1 | NonReactive Control Results Q1.P1 | Standard Acceptance Criteria NonReactive Control Results Q1.P1 Equal to Acceptable
RPR Daily | RPR Daily P1 | Weak Reactive Control Results Q4.P1 | Standard Acceptance Criteria Weak Reactive Control Results Q4.P1 Equal to Acceptable
RPR Daily | RPR Daily P1 | Reactive Control Results Q2.P1 | Standard Acceptance Criteria Reactive Control Results Q2.P1 Equal to Acceptable
RPR Daily | RPR Daily P1 | Problem with QC Q3.P1 | Standard Acceptance Criteria Problem with QC Q3.P1 Equal to No